## INDEX
Links to this File's Tables and Notes
Table YE1 | Definitions for the Different Interface Points
Table YE2 | Yield Enhancement Difficult Challenges
Table YE3 | Technology Requirements for Wafer Environmental Contamination Control
Table YE4 | AMC Monitoring Methods
Table YE4a | Supporting Table for On-line Methods
Table YE5 | Yield requirements for MEMS production
Table Notes
Table YE3 notes
Table YE4 notes
Table YE4a notes
Table YE5 notes
Potential Solutions Figures
Figure YE5 | Wafer Environment and Reticle Environment Contamination Measurement and Control Potential Solutions
Figure YE6 | Substrate Environment Contamination Control Potential Solutions - Critical Components
Figure YE7 | Substrate Environment Contamination Control Potential Solutions - UPW
Figure YE8 | Substrate Environment Contamination Control Potential Solutions - Liquid Chemicals
Copyright © 2018 IEEE. All rights reserved.

## Table YE1
INDEX
Table YE1 Definitions for the Different Interface Points
|  | POS | POD | POC | POE | POU | POP
|  | Delivery Point of Gas/Chemical Supplier | Outlet of Central Facility System | Submain or VMB/VMP Take off Valve | Entry to Equipment or Sub Equipment | Entry to the Process Chamber | Contact with Wafer
| Ultrapure water | Raw water | Outlet of final filtration in UPW plant | Outlet of submain take off valve | Inlet of wet bench or subequipment | Inlet of wet bench bath, spray nozzle, or connection point to piping, which is also used for other chemicals | Wafer in production
| Process chemicals | Chemical drum/tote/bulk supply | Outlet of final filtration of chemical distribution unit | Outlet of VMB valve | Inlet of wet bench or intermediate tank | Inlet of wet bench bath or spray nozzle | Wafer in production
| Specialty gases | Gas cylinder or bulk specialty gas systems | Outlet of final filtration of gas cabinet | Outlet of VMB valve | Inlet of equipment | Inlet of chamber (outlet of MFC) | Wafer in production
| Bulk gases | Bulk gas delivered on site or gas generator | Outlet of final filtration/purification | Outlet of submain take off valve or VMB valve | Inlet of equipment/ subequipment | Inlet of chamber (outlet of MFC) | Wafer in production
| Cleanroom and AMC | Outside air | Outlet of make-up air handling unit | Outlet of filters in cleanroom ceiling | Inlet to mini-environment or sub equipment for AMC, outlet of the tool filter for particles | Gas/air in vicinity to wafer/substrate | Wafer/substrate in production (AMC/ SMC)

## Table YE2-2017
INDEX | Table YE2 Yield Enhancement Difficult Challenges – 2017 SECC
| Near-Term Challenges: 2017-2024 | Description
| Challenge #1 Control of Electrically Active and other particles in Ultrapure Water and Liquid Chemicals for advanced semiconductor manufacturing (logic key driver) | Insufficient metrology and marginal treatment technology capability
| Challenge #2 Control of Electrically Active and other particles in UPW of EUV mask production | Insufficient metrology and marginal treatment technology capability
| Challenge #3 Critical Organics in UPW | Insuffcient metrology for online speciation of organics
| Challenge #4 Metals in H2O2 | Insufficient metrology capability. Characteriztion needed for specific critical metals for key process steps and segments.
| Challenge #5 Slurry particle characterization and other impurity monitoring | Insufficient metrology capability
| Challenge #6 Plating chemicals: particles > critical size (>0.150 µm) (#/ml) | Need reliable monitoring of large particles
| Challenge #7 Liquid Chemicals: particles > critical size (#/ml) | Need reliable monitoring of particles <10 nm. Need Filtration technology that is relaible at the stated retention rating and that has minimal "rise up" contributions at the time of install into the process.
| Challenge #8 CMP Slurries: large particles > MPS (#/ml) | Need reliable LPC metrology that can measure targeted particles sizes and concentrations that ignore the MPS sizes and do not require sample dilution to obtain the measurement.
| Challenge #9 Post-CMP (pCMP) Cleans particles | Some pCMP cleans will have volatile or highly reactive ingredients that form bubbles in solution, therefore; new metrology and new techniques will be needed to differentiate between bubbles and real particles for contamination control. The pCMP cleans that contain surfactants will have micelle particles and structures that are necessary for the chemistry to perfrom the intended function. The need is for reliable and accurate metrology to distinguish micelles from other foreign particles in solution.
| Challenge #10 Particle measurement for ISO class 1 | Need more efficient and effective laser particle counters with detection limit << 100 nm or CPC with sample flow >> 0,1 cfm
| Challenge #11 AMC metrology | Need of more sensitive gas analyzers for HCl, HF, HNO3, HBr, NH3, Total Acids, H2S, SO2, organic acids, sulphuric acids and Total bases with detection limit << 100 ppt
| Challenge #12 Particle contribution by critical components and its control | Insufficient analytical capability and unknown level of cleanliness
| Long-Term Challenges: 2025–2032 | Description
OUT OF CURRENT SCOPE OF YIELD ENHANCEMENT

## Table YE3-2017
INDEX | Table YE3 Technology Requirements for Wafer Environmental Contamination Control
| Year of Production | 2017 | 2018 | 2019 | 2020 | 2021 | Remarks | Driver, actions required
| Logic industry "Node Range" Labeling (nm) | "10" | "10" | "7" | "7" | "5"
| IDM-Foundry node labeling | i10-f7 | i10-f7 | i7-f5 | i7-f5 | i5-f3
| Logic device structure options | finFET FDSOI | finFET FDSOI | finFET LGAA | finFET LGAA | LGAA VGAA
| Logic device mainstream device | finFET | finFET | finFET | finFET | LGAA
| MPU/SoC Metalx ½ Pitch (nm)[1,2] | 18 | 18 | 14 | 14 | 12 | based on More Moore table
| Critical particle size non-electrically active (non-EAP) (nm) based on Logic 1/2 Pitch (nm (contacted) [1] | 18 | 18 | 14 | 14 | 12
| Critical particle size Eelctrically Active (nm) based on 50% of Logic 1/2 Pitch (nm (contacted) [1] | 9 | 9 | 7 | 7 | 6
Wafer Environment Control such as Cleanroom, SMIF POD, FOUP, etc….not necessarily the cleanroom itself but wafer environment.
| Number of particles (/m3) [1] [2] | ISO CL1 | ISO CL1 | ISO CL1 | ISO CL1 | ISO CL1
Airborne Molecular Contaminants in Gas Phase (pptV, V for Volume)) [3, 7, 12,13,14,15,33].
Lithography: POI (wafer stage & retickle library) [23]
| Total Inorganic Acids | 5000 | 5000 | 5000 | 5000 | 5000 | Considering Exposure conditions for lenses & reticles in combination with the 99% efficiency of the tool chemical filtration system | lens & mask hazing
| Total Organic Acids [30] | 2000 | 2000 | tbd | tbd | tbd | Pic et al, Proceedings of SPIE Vol. 7638, 76380M (2010). | lens hazing
| Total Bases | 20000 | 20000 | 50000 | 50000 | 50000 | Considering Exposure conditions for lenses & reticles in combination with the 99% efficiency of the tool chemical filtration system | Chemically amplified resist reaction, hazing
| Volatile Organics (w/ GCMS retention times ≥ benzene, calibrated to hexadecane) [31] | 26000 | 26000 | 26000 | 26000 | 26000 | Considering Exposure conditions for reticles in combination with the 99% efficiency of the tool chemical filtration system | hazing
| Refractory compounds (organics containing for example S, P, Si) [40] [43] | 2000 | 2000 | 2000 | 2000 | 2000 | POE Chemical filtration system, considering low efficiency of filtration system | hazing, SiO2 deposition
| Lithography POI (inspection tools stage) | [6 empty] | Considering a 90% efficiency of chemical filtration systems to tracks & inspection at POE
| Total Inorganic Acids | 2000 | 2000 | 2000 | 2000 | 2000 | Fontaine et al., Photomask Technology 2010, Paper SPIE 7823-61 1) | 1) Considering surface contamination coverage on
| Total Organic Acids [30] | 2000 | 2000 | 2000 | 2000 | 2000 | Fontaine et al, Proc. SPIE, Vol. 7545 I, I1-I12 (2010); Fontaine et al, CEA-Leti (2010), shared information 2011 | pod equilibrium with clean room within 10-20 hours
| Total Bases | 2000 | 2000 | 2000 | 2000 | 2000 | Fontaine et al., Photomask Technology 2010, Paper SPIE 7823-61 1) | and considering a pod residence time in zone of 1-2 hours
| Condensable organics (definition to SEMI F21-95, bp 150°C) | 1000 | 1000 | 1000 | 1000 | 1000 | Fontaine et al, Proc. SPIE, Vol. 7545 I, I1-I12 (2010): Fontaine et al, CEA-Leti (2010), shared information 2011
| Refractory compounds (organics containing for example S, P, Si) | tbd | tbd | tbd | tbd | tbd | valid for inspection tools only
| Reticle Storage (inside stocker, inside pod, inside exposure tool library, inside inspection tool) | [6 empty] | XCDA to BRS may require more stringent contamination limits
| Sum of Inorganic Acids (HCl, HF, HNO3) | <200 | <200 | <200 | <200 | <200 | Fontaine et al., Photomask Technology 2010, Paper SPIE 7823-61; reference Cr mask | haze/particle formation from surface deposited contamination
| Total Organic Acids [30] | <200 | <200 | <200 | <200 | <200 | Fontaine et al, Proc. SPIE, Vol. 7545 I, I1-I12 (2010); Fontaine et al, CEA-Leti (2010), shared information 2011 | haze/particle formation from surface deposited contamination
| Total Bases | <200 | <200 | <200 | <200 | <200 | Fontaine et al., Photomask Technology 2010, Paper SPIE 7823-61; reference Cr mask | haze/particle formation from surface deposited contamination
| Condensable organics (definition to SEMI F21-95, bp 150°C) | <100 | <100 | <100 | <100 | <100 | Fontaine et al, Proc. SPIE, Vol. 7545 I, I1-I12 (2010); Fontaine et al, CEA-Leti (2010), shared information 2011 | carbon particle formation from surface deposited contamination
| Refractory Compounds | tbd | tbd | tbd | tbd | tbd |  | valid for inspection tools
Gate/Furnace area wafer environment (FOUP inside)
| Total metals [8] (E+10 atoms/cm2/day) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | Benchmarking | CV (mobile ions, e.g. Na); Lifetime (Ca);V(t) (Fe);GOI (Ca);TDDB
| Dopants [4] (E+10 front end of line only) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | Benchmarking | Resistivity
| Volatile Organics (w/ GCMS retention times ≥ benzene, calibrated to hexadecane) [31] | 2000 | 2000 | 2000 | 2000 | 2000 | Value assumed identical to FOUP ambient due to lack of data
| SMC (surface molecular condensable) organics on wafers, ng/cm2/day [12] | tbd | tbd | tbd | tbd | tbd | Literature Sony | Effect on DOP, BHT, Review value & definition (e.g. define critical contaminants separetely)
| Total SM (surface metals) on wafer, E+10 atoms/cm2/day | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | from FEP surface preparation table
Salicidation Wafer Environment (FOUP inside; wafer environment)
| Total Inorganic Acids [44] | 2000 | 2000 | 2000 | 2000 | 2000 | feedback IDM, NOx is excluded; may be specific for type of silicide | Corrosion
| Total Organic Acids [30] | 5000 | 5000 | 5000 | 5000 | 5000 | feedback IDM, NOx is excluded; may be specific for type of silicide | Corrosion
| Exposed Copper (POI, Wafer environment) | [5 empty] | Values do not consider AMC filtration at the inlet of the tools. If filtration is used consider the efficiency of your system. | without Chemical filtration at POE to tool
| HCl | tbd | tbd | tbd | tbd | tbd
| HF | tbd | tbd | tbd | tbd | tbd
| HBr | tbd | tbd | tbd | tbd | tbd
| HNO3 | tbd | tbd | tbd | tbd | tbd
| Total Bases | 2000 | 2000 | 2000 | 2000 | 2000 | Bases can create corrosion through particle deposition e.g. (NH4)2SO4. Value to be further discussed.
| Total Organic Acids [30] | 500 | 500 | 500 | 500 | 500 | More discussion required (one data point)
| Total other corrosive species [32] | 1000 | 1000 | 1000 | 1000 | 1000 | Benchmarking | Filter capacity
| H2S | 1000 | 1000 | 1000 | 1000 | 1000 | Benchmarking | Filter capacity
| Total sulphur compounds | 2500 | 2500 | 2500 | 2500 | 2500 | Benchmarking | Filter capacity
| Exposed Copper Wafer Environment (FOUP inside) | [5 empty] | Defined for standard cleanroom r.h., for other r.h. refer to literature. Based on 24h exposure. | for 24 hours sit time
| HCl | 200 | 200 | 200 | 200 | 200 | Fontaine et al., H53ECS Transactions, 25 (5) 79-86 (2009), < 1E13 atoms/cm2 | Corrosion
| HF | 5000 | 5000 | 5000 | 5000 | 5000 | Fontaine et al., ECS Transactions, 25 (5) 79-86 (2009), < 1E14 atoms/cm2 | Corrosion
| HBr | tbd | tbd | tbd | tbd | tbd |  | Corrosion
| HNO3 | tbd | tbd | tbd | tbd | tbd |  | Corrosion
| Total Organic Acids [30] | 100 | 100 | 100 | 100 | 100 | Benchmarking | Corrosion
| Total bases | tbd | tbd | tbd | tbd | tbd | IDM ? | Corrosion
| Total other corrosive species [32] | tbd | tbd | tbd | tbd | tbd | Benchmarking | Corrosion
| H2S | tbd | tbd | tbd | tbd | tbd | Benchmarking | Corrosion
| Total sulphur compounds | 5000 | 5000 | 5000 | 5000 | 5000 | Benchmarking | Corrosion
| Moisture (ppb) | tbd | tbd | tbd | tbd | tbd | Benchmarking | Corrosion
| Exposed Aluminum Wafer Environment (FOUP inside) | [5 empty] | Defined for standard cleanroom r.h., for other r.h. refer to literature. Based on 24h exposure.
| Total Inorganic Acids | tbd | tbd | tbd | tbd | tbd | Benchmarking | Corrosion
| HCl | 100 | 100 | 100 | 100 | 100 | From Bi Jun Wu et Al. IEEE trans. On semicond. Manuf., 23 (4) 2010, and calculation from H Fontaine, < 8E12 atoms/cm2 for 24 hrs | Corrosion
| HF | 200 | 200 | 200 | 200 | 200 |  | Corrosion
| HBr | tbd | tbd | tbd | tbd | tbd |  | Corrosion
| HNO3 | tbd | tbd | tbd | tbd | tbd |  | Corrosion
| Total Organic Acids [30] | tbd | tbd | tbd | tbd | tbd | Discussion | Corrosion
| Total other corrosive species [32] | tbd | tbd | tbd | tbd | tbd | Benchmarking | Corrosion
Critical areas (Litho, Metrology)
| Temperature range in +/-K at POE [37] | 1 | 1 | 1 | 1 | 1 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Maximum short term temperature variation at POE in +/-K/5 min [37] | 1 | 1 | 1 | 1 | 1 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Maximum long term temperature variation in +/-K/hour at POE [37] | 1 | 1 | 1 | 1 | 1 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Humidity range in +/- % relative humdity r.H. at POE [37] | 3 | 3 | 3 | 3 | 3 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Humidity range in +/- % relative humidity r.H. at POE [37] | 3 | 3 | 3 | 3 | 3 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Maximum short term humidity variation in +/-r.H./5 min at POE [37] | 2 | 2 | 2 | 2 | 2 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
Non-critical areas (others than Litho, Metrology)
| Temperature range on +/-K at POE [37] | 2 | 2 | 2 | 2 | 2 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Humidity range in +/- % relative humdity r.H. at POE [37] | 5 | 5 | 5 | 5 | 5 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
| Humidity range in +/- % relative humidity r.H. at POE [37] | 5 | 5 | 5 | 5 | 5 | Benchmarking, Discussion with Factory integration TWG | Inlet spec Environmental chamber, Air density
Ultrapure Water [29]
| Total organic carbon (ppb) [22] for immersion litho, POE | <1 | <1 | <1 | <1 | <1 |  | lense hazing
| Critical Organics as C (ppb) [22], POE | <1 | <1 | <1 | <1 | <1 |  | 1E+12 at/cm2 of C on the wafer
| Non-critical organics as C (ppb) [22], POE | <3 | <3 | <3 | <3 | <3 |  | typical UPW system capability
| Reactive silica (ppb) as SiO2 [18], POC | <0.05 | <0.05 | <0.05 | <0.05 | <0.05
| Number of EAP at critical particle size (#/L) [26], POE | 10000 | 10000 | 10000 | 10000 | 10000 | EAP - electrically active particles | electrical conductivity may be lower for very small particles - needs further investigation
| Number of non-EAP at critical size (#/L) [26], POE | 10000 | 10000 | 10000 | 10000 | 10000
| Number of particles for EUV mask production (#/L) [26.3], POE | 1000 | 1000 | 1000 | 1000 | 1000 | An order of magnitude stricter than FAB
| Critical size (EUV mask), nm | 12 | 12 | 12 | 12 | 12 | function of the wavelenght
| Dissolved oxygen (ppb) (contaminant based) [16] POE | <10 | <10 | <10 | <10 | <10
| Dissolved nitrogen (ppm) [10] POE | 8–18 | 8–18 | 8–18 | 8–18 | 8–18
| Hydrogen Peroxide (ppb) [42] POD | <10 | <10 | <10 | <10 | <10 | consistent with DO and supported by experiment
| Metals (ppt each) (Al, As, Ba, Ca, Cd, Co, Cr, Cu, Fe, K, Li, Mg, Mn, Mo, Na, Ni, Pb, Sb, Sn, Sr, Ti, V, W, Zn + Pt*) [39], POP | <1 | <1 | <1 | <1 | <1
| Critical Metals for image sensors (ppt each) (Cu, Mn) [39], POP | <0.2 | <0.2 | <0.2 | <0.2 | <0.2 | based on published experiemental data; | did not take into account thermal processes, which may require even tighter limit for these and some other metals to be considered due to their diffusivity. Further investigation is required.
| Other critical ions (ppt each) [24], POP | <50 | <50 | <50 | <50 | <50
| Phosphate as HPO4 (ppt), POP | <20 | <20 | <20 | <20 | <20
| Boron (ppt) [24], POP | <50 | <50 | <50 | <50 | <50
| Temperature stability (K) [38], POE | ± 1 | ± 1 | ± 1 | ± 1 | ± 1
| Temperature gradient in K/10 [38] for immersion litho, POE [38] | <0.1 | <0.1 | <0.1 | <0.1 | <0.1
Liquid Chemicals
| Cleaning chemistries (aqueous and solvent): Number of EAP particles at critical particle size (#/L) [26], POE | 10000 | 10000 | 10000 | 10000 | 10000 | EAP - electrically active particles | electrical conductivity may be lower for very small particles - needs further investigation
| Cleaning chemistries (aqueous and solvent): Critical size of EAP, nm | 12 | 12 | 12 | 12 | 12
| Cleaning chemistries (aqueous and solvent): Number of all particles at critical size (#/L) [26], POE | 10000 | 10000 | 10000 | 10000 | 10000
| 49% HF: Na, K, Fe, Ni, Cu, Cr, Co, Ca, (Ag, Au, Pd, Pt, Ru, Mo) (ppt, each) [21] | 150 | 150 | 150 | 150 | 150
|  | 100 | 100 | 100 | 100 | 100 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 49% HF: all other metals not listed in row above (ppt, each) [20] | 500 | 500 | 500 | 500 | 500 | No change
| 49% HF: Anions (ppb) | 500 | 500 | 500 | 500 | 500 | No change. Needs a well defined method and metrology standard.
| 49% HF: Cl (ppb) | 1000 | 1000 | 1000 | 1000 | 1000
|  | 100 | 100 | 100 | 100 | 100 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 29% NH4OH: Al, Na, K, Fe, Ni, Cu, Cr, Co, Ca, (Au, Ba, Cd, Mg, Mn, Mo, Pb, Pd, Pt, Ru, Sn, Ti, V, W, Zn) (ppt, each) [21] | 150 | 150 | 150 | 150 | 150
|  | 100 | 100 | 100 | 100 | 100 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 29% NH4OH: all other metals not listed in row above (ppt, each) [21] | 500 | 500 | 500 | 500 | 500 | No change
| 29% NH4OH: Anions (ppb) | 10 | 10 | 10 | 10 | 10 | No change. Needs a well defined method and metrology standard.
| 100% IPA: Al, Ba, B, K, Na, K, Fe, Li, Mg, Mn, Pb, Sn, Ti, Zn, Ni, Cu, Cr, Co, Ca (ppt, each) [28] | 150 | 150 | 150 | 150 | 150
|  | 100 | 100 | 100 | 100 | 100 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 100% IPA: all other metals not listed in row above (ppt, each) [21] | 500 | 500 | 500 | 500 | 500 | No change
| 100% IPA: High molecular weight organics (ppb) | 500 | 500 | 500 | 500 | 500 | No change. Needs a well defined method and metrology standard.
| 100% IPA: Anions (ppb) | 10 | 10 | 10 | 10 | 10 | No change. Needs a well defined method and metrology standard.
| 30% H2O2: Al, Na, K, Fe, Ni, Cu, Cr, Co, Ca, (Ag, Au, Ba, Cd, Mg, Mn, Mo, Pb, Pd, Pt, Ru, Sn, Ti, V, W, Zn) (ppt, each) [21] | 150 | 150 | 150 | 150 | 150
|  | 10 | 10 | 10 | 10 | 10 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 30% H2O2: All other metals not listed in row above (ppt, each) [21] | 500 | 500 | 500 | 500 | 500 | No change
| 30% H2O2: Anions (ppb) | 10 | 10 | 10 | 10 | 10 | No change. Needs a well defined method and metrology standard.
| 30% H2O2: resin byproducts (ppb) e.g. total amines [41] | 5 | 5 | 5 | 5 | 5 | No change. Needs a well defined method and metrology standard.
| 37% HCL: Ba, Cu, K, Pb, Li, Mn, Ni, Cu, Cr, Co, (ppt, each) | 1000 | 1000 | 1000 | 1000 | 1000
|  | 100 | 100 | 100 | 100 | 100 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 37% HCl: all other metals not listed in row above (ppt, each) [20] | 10000 | 10000 | 10000 | 10000 | 10000 | No change
| 37% HCl: Anions (ppb) | 300 | 300 | 300 | 300 | 300 | No change. Needs a well defined method and metrology standard.
| 96% H2SO4: Al, Ba, B, Ca, K, Pb, Li, Mg, Mn, Na, Sn, Ti, Zn, K, Ni, Cu, Cr, Co, (ppt, each) | 1000 | 1000 | 1000 | 1000 | 1000
|  | 100 | 100 | 100 | 100 | 100 | Values based on reported needs of the industry and incorporated in the most recent version of the associated SEMI standard.
| 96% H2SO4: all other metals not listed in row above (ppt, each) [20] | 10000 | 10000 | 10000 | 10000 | 10000 | No change
| 96% H2SO4: Anions (ppb) | 200 | 200 | 200 | 200 | 200 | No change. Needs a well defined method and metrology standard.
| CMP slurries: scratching particles (#/ml >2x Mean Particle Size) [9] [17] | TBD | TBD | TBD | TBD | TBD
|  | 100 | 100 | 100 | 100 | 100 | For Mullti-Modal abrasives, values are based on the MPS of the largest Mode. This is to measure the amount agglomeration that cause scratch defects during the CMP process.
| Post-CMP clean chemicals: particles >0.065 (/ml) [1] [9] [17] | TBD | TBD | TBD | TBD | TBD
| Post-CMP clean chemicals: particles >0.065 µm (/ml) [1] [9] [17] [49] Non-Surfactant Based | 500 | 400 | 200 | 100 | 50 | Chemistry containing Micelles will require a larger particle size threshhold to be used to prevent counting Micelles as harmfull particles.
| Post-CMP clean chemicals: particles >0.2 µm (/ml) [1] [9] [17] [49] Surfactant Based | 100 | 90 | 80 | 70 | 60 | Chemistry containing Micelles will require a larger particle size threshhold to be used to prevent counting Micelles as harmfull particles.
| Plating chemicals: particles > critical size (>0.150 µm) (#/ml) [1] [9] [17] [50] | TBD | TBD | TBD | TBD | TBD
|  | 1000 | 1000 | 1000 | 1000 | 1000 | Chemistry containing Micelles will require a larger particle size threshhold to be used to prevent counting Micelles as harmfull particles.
Bulk Gases (Contaminants, ppbv) [5]
| N2 (O2, H2, H2O, CO, CO2, ≤C4, ≥C5) [34] | < 1 | < 1 | < 1 | < 1 | < 1 | Post purification spec
| O2 (N2) | <50 | <50 | <50 | <50 | <50
| O2 (Ar) | <500 | <500 | <500 | <500 | <500 | As delivered when required.
| O2 (H2, H2O, CO, CO2, ≤C4, ≥C5) | [5 empty] | no change to outlet spec
| Ar (O2, H2, H2O, CO, CO2, ≤C4, ≥C5) [34] | <1 | <1 | <1 | <1 | <1 | Post purification spec
| H2 (N2) | <1 | <1 | <1 | <1 | <1 | Post purification spec, when required
| H2 (O2, H2O, CO, CO2, ≤C4, ≥C5) | <1 | <1 | <1 | <1 | <1 | Post purification spec
| He (O2, H2, H2O, CO, CO2, ≤C4, ≥C5) [34] | <1 | <1 | <1 | <1 | <1 | Post purification spec
| CO2 (CO, H2O, O2, ≤C4, ≥C5) | TBD | TBD | TBD | TBD | TBD | TBD
Lithography and Critical Metrology Purge Gases (ppbv)
| Critical clean dry air (H2O) | <1000 | <1000 | <1000 | <1000 | <1000
| Critical clean dry air (H2, CO) | <2000 | <2000 | <2000 | <2000 | <2000
| Critical clean dry air (organics (molecular weight > benzene) normalized to hexadecane equivalent) | <3 | <3 | <3 | <3 | <3
| Critical clean dry air (total base as NH3) | <1 | <1 | <1 | <1 | <1
| Critical clean dry air (NH3 (as NH3)) | <0.1 | <0.1 | <0.1 | <0.1 | <0.1
| Critical clean dry air (total acid including SO2 (as SO4)) | <0.1 | <0.1 | <0.1 | <0.1 | <0.1
| Critical clean dry air (SO4 (as SO4)) | <0.03 | <0.03 | <0.03 | <0.03 | <0.03
| Critical clean dry air (Each refractory compound (0rganics containing S, P, Si) | <0.1 | <0.1 | <0.1 | <0.1 | <0.1
| Lithography nitrogen tool/maintenance purging gas supply (H2O, O2, CO2) | <500 | <500 | <500 | <500 | <500
| Lithography nitrogen tool/maintenance purging gas supply (CO) | <2000 | <2000 | <2000 | <2000 | <2000
| Lithography nitrogen tool/maintenance purging gas supply (H2) | <2000 | <2000 | <2000 | <2000 | <2000
| Lithography nitrogen tool/maintenance purging gas supply (organics (molecular weight > benzene) normalized to hexadecane equivalent) | <3 | <3 | <3 | <3 | <3
| Lithography nitrogen tool/maintenance purging gas supply (total base (as NH3)) | <0.15 | <0.15 | <0.15 | <0.15 | <0.15
| Lithography nitrogen tool/maintenance purging gas supply (total acid (as SO4) including SO2) | <0.025 | <0.025 | <0.025 | <0.025 | <0.025
| Lithography nitrogen tool/maintenance purging gas supply (refractory compounds (organics containing S, P, Si, etc.) normalized to hexadecane equivalent) | <0.1 | <0.1 | <0.1 | <0.1 | <0.1
| Lithography helium tool/maintenance purging gas supply (H2O) | <3500 | <3500 | <3500 | <3500 | <3500
| Lithography helium tool/maintenance purging gas supply (O2, CO2) | <4000 | <4000 | <4000 | <4000 | <4000
| Lithography helium tool/maintenance purging gas supply (CO, H2) | <10000 | <10000 | <10000 | <10000 | <10000
| Lithography helium tool/maintenance purging gas supply (organics(molecular weight > benzene) normalized to hexadecane equivalent) | <3 | <3 | <3 | <3 | <3
| Lithography helium tool/maintenance purging gas supply (total base (as NH3)) | <0.15 | <0.15 | <0.15 | <0.15 | <0.15
| Lithography helium tool/maintenance purging gas supply (total acid including SO2 (as SO4)) | <0.025 | <0.025 | <0.025 | <0.025 | <0.025
| Lithography helium tool/maintenance purging gas supply (refractory compounds (organics containing S, P, Si, etc.) normalized to hexadecane equivalent) | <0.1 | <0.1 | <0.1 | <0.1 | <0.1
| Number of particles > critical size (/M3) [1] | 100 | 100 | 100 | 100 | 100
| Number of particles > 0.1 µm (/M3) [1] | 0.20004 | 0.14429 | 0.1 | 0.070497 | 0.0512
Etchants (Corrosive, e.g., BCl3, Cl2, HBr)
| O2, H2O (ppbv) | <500 | <500 | <500 | <500 | <500
| Critical specified metals/total metals (ppbw) [19] | <1/1000 | <1/1000 | <1/1000 | <1/1000 | <1/1000
Etchants (Non-corrosive, e.g., C5F8, C4F8, C4F6, CH2F2, SF6)
| O2, H2O (ppbv) | <1000 | <1000 | <1000 | <1000 | <1000
Etchants (e.g. Xe)
| O2, H2O (ppbv) | <1000 | <1000 | <1000 | <1000 | <1000
ALD/CVD blanket / purge/ delivery gases
| Sweep and bubbler gases for ALD/CVD volatile chemical delivery e.g., N2 (O2, H2, H2O, CO, CO2, THC) (ppb) [34] | tbd | tbd | tbd | tbd | tbd
CVD/ALD Thin Film Process Fluids
CVD/ALD Thin Film Process Fluids - High k Dielectrics
High k Gate Dielectrics e.g., Al2O3 and LaAlO3
| High k Gate Dielectrics - TMA [75-24-1] (Assay) | 99.5 | 99.5 | 99.5 | 99.5 | 99.5
| High k Gate Dielectrics - TMA (O2 ppm) | <10 | <10 | <10 | <10 | <10
| High k Gate Dielectrics - TMA Metals each element (ppb) | <150 | <150 | <150 | <150 | <150
| High k Gate Dielectrics - TMA (Silicon ppm) | <1 | <1 | <1 | <1 | <1
| High k Gate Dielectrics - TMA (Hydrocarbons ppm) | <5 | <5 | <5 | <5 | <5
| High k Gate Dielectrics - La(FMD)3 (Lanthanum tris(N,N’-isopropylformamidinate) (assay) | tbd | tbd | tbd | tbd | tbd
High k Gate Dielectrics - Hf oxides, silicates or aluminates
| High k Gate Dielectrics - Hf- HTB [2172-02-3] (assay) | tbd | tbd | tbd | tbd | tbd
| High k Gate Dielectrics - Hf- HfCl4 [13499-05-3] (assay) | tbd | tbd | tbd | tbd | tbd
High k Gate Dielectrics - Si source - Metal Silicates
| Si source - Metal Silicates - 3DMAS [15112-89-7] (assay) | tbd | tbd | tbd | tbd | tbd
| Si source - Metal Silicates - 4DMAS [1624-01-7] (assay) | tbd | tbd | tbd | tbd | tbd
Va High k Gate Dielectrics - MgO
| High k Gate Dielectrics - MgO - Mg(EtCp)2 [114460-02-5] (assay) | tbd | tbd | tbd | tbd | tbd
High k Capacitor Dielectrics e.g., Al2O3 and LaAlO3
High k Capacitor Dielectrics - TMA [75-24-1] (see High k Gate Dielectrics - TMA above)
| High k Capacitor Dielectrics - TDEAA [32093-39-3 ] (assay) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - La(thd)3 [14319-13-2] (assay) | tbd | tbd | tbd | tbd | tbd
High k Capacitor Dielectrics - Ta2O5
| High k Capacitor Dielectrics - Ta(OC2H5)5 [6074-84-6] (assay) | tbd | tbd | tbd | tbd | tbd
High k Capacitor Dielectrics - Hf oxides
| High k Capacitor Dielectrics - HfCl4 [13499-05-3] (assay) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - HfCl4 (Zr -ppm) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - HfCl4 (Metals each element -ppb) | tbd | tbd | tbd | tbd | tbd | semi has 200 to 500 ppb for most but Ti/Al higher
| High k Capacitor Dielectrics - HfCl4 (Hydrocarbons -ppm) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics HfCl4 (Hydroloysis product Hf oxides -ppm) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics Hf metal organics - TDEAH [19824-55-6 ] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TDEAH (Zr -ppm) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TDEAH (Metals each element -ppb) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TDEAH (Hydrocarbons -% by NMR) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TDEAH (Chloride ppm) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TDEAH (Hydroloysis product alkylamine -ppm) | tbd | tbd | tbd | tbd | tbd
High k Capacitor Dielectrics - Hf metal organics, TEMAH [352535-01-4 ] (same as TDEAH )
| High k Capacitor Dielectrics - Zr metal organics - TEMAZ [175923-04-3] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TEMAZ - (Hf ppm) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TEMAZ (Metals each element ppb) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TEMAZ (Hydrocarbons (% by NMR) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - TEMAZ (Chloride ppm) | tbd | tbd | tbd | tbd | tbd | available at this level
| High k Capacitor Dielectrics - TEMAZ (Hydroloysis product (alkylamine) ppm ) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - SrTiO3 | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - Sr(DPM)2 (assay) | tbd | tbd | tbd | tbd | tbd | fix this
| High k Capacitor Dielectrics - Ti(OiPA)4 [546-68-9] (assay) | tbd | tbd | tbd | tbd | tbd
| High k Capacitor Dielectrics - Ti(O)dpm)2 [152248-67-4] (assay) | tbd | tbd | tbd | tbd | tbd
Spacers; misc. FEOL dielectric Films
Spacer Dielectric - SiCN
| Spacer Dielectric - BTBAS [186598-40-3] (assay) | tbd | tbd | tbd | tbd | tbd
CVD/ALD Thin Film Process Fluids - Low k Dielectrics
Interconnect Dielectrics - Barrier / Etch Stops
Barrier / Etch Stops - SiC
| Barrier / Etch Stops - DMDMOS [1112-39-6] (assay) | tbd | tbd | tbd | tbd | tbd
| Barrier / Etch Stops - 3MS [993-07-7] (Assay) | tbd | tbd | tbd | tbd | tbd
| Barrier / Etch Stops - 4MS [75-76-3] (assay) | tbd | tbd | tbd | tbd | tbd
| Barrier / Etch Stops SiN / SiO2 | tbd | tbd | tbd | tbd | tbd
| Barrier / Etch Stops SiN / SiO2 - SiH4 [7803-62-5] (assay) | tbd | tbd | tbd | tbd | tbd
| Barrier / Etch Stops SiN / SiO2 SiCN | tbd | tbd | tbd | tbd | tbd
Barrier / Etch Stops SiN / SiO2 - 4MS [75-76-3] (assay)
| Barrier / Etch Stops SiN / SiO2 - 3MS [993-07-7] (assay) | tbd | tbd | tbd | tbd | tbd
Barrier / Etch Stops - SiCO
Barrier / Etch Stops - SiCO - Alkylsilanes (e.g., Trimethylsilane [993-07-7] , Tetramethylsilane [75-76-3], DMDMOS [1112-39-6], OMCTS [556-67-2]) [25] - ILD
| SiCO - Alkylsilanes (Metals except B, Au, Ag (ppb, each)) | tbd | tbd | tbd | tbd | tbd
| SiCO - Alkylsilanes (B, Au, Ag (ppb, each)) | tbd | tbd | tbd | tbd | tbd
| SiCO - Alkylsilanes (H2O (ppm)) | tbd | tbd | tbd | tbd | tbd
| SiCO - Alkylsilanes (CO, CO2 (ppm)) | <0.5 | <0.5 | <0.5 | <0.5 | <0.5
| SiCO - Alkylsilanes (Non-methane hydrocarbons C2-C4 (ppm)) | <4 | <4 | <4 | <4 | <4
| SiCO - Alkylsilanes (Nitrogen (ppm)) | <2 | <2 | <2 | <2 | <2
| SiCO - Alkylsilanes (Ar+O2 (ppm)) | <0.5 | <0.5 | <0.5 | <0.5 | <0.5
| SiCO - Alkylsilanes (Chloride (ppm)) | <2 | <2 | <2 | <2 | <2
| SiCO - Alkylsilanes - 3MS [993-07-7] (assay) | tbd | tbd | tbd | tbd | tbd
| SiCO - Alkylsilanes - DEMS [2031-62-1] (assay) | tbd | tbd | tbd | tbd | tbd
| SiCO - Alkylsilanes - OMCTS [556-67-2] (assay) | tbd | tbd | tbd | tbd | tbd
| SiCO - Alkylsilanes - DMDMOS [1112-39-6] (assay) | tbd | tbd | tbd | tbd | tbd
| Interconnect Bulk Low k dielectrics | tbd | tbd | tbd | tbd | tbd
| Interconnect Bulk Low k dielectrics - FSG / FOSG | tbd | tbd | tbd | tbd | tbd
| FSG / FOSG - SiF4 [7783-61-1 ] (assay) | tbd | tbd | tbd | tbd | tbd
| FSG / FOSG - TEOS [78-10-4] (assay) | tbd | tbd | tbd | tbd | tbd
| FSG / FOSG - SiH4 [7803-62-5] (assay) | tbd | tbd | tbd | tbd | tbd
| FSG / FOSG - 3MS [993-07-7] (assay) | tbd | tbd | tbd | tbd | tbd
Interconnect Bulk Low k dielectrics - OSG / SiCO (k = 3.0 to k = 2.7)
OSG / SiCO - 3MS [993-07-7] (assay)
| OSG / SiCO - DEMS [2031-62-1] (assay) | tbd | tbd | tbd | tbd | tbd
| OSG / SiCO - TMCTS [2370-88-9] (assay) | tbd | tbd | tbd | tbd | tbd
| OSG / SiCO - DMDMOS [1112-39-6] (assay) | tbd | tbd | tbd | tbd | tbd
| OSG / SiCO - OMCTS [556-67-2] (assay) | tbd | tbd | tbd | tbd | tbd
Interconnect Ultra low k Porous OSG / SiCO (k < 2.7 to 2.2)
| Ultra low k Porous OSG / SiCO - DEMS [2031-62-1] (assay) | tbd | tbd | tbd | tbd | tbd
| Ultra low k Porous OSG / SiCO - DEMS ( Organic Porogen (%)) | tbd | tbd | tbd | tbd | tbd
CVD/ALD Thin Film Process Fluids - Metal Gate Electrodes
Metal Gate Electrodes - Titanium nitride/carbonitride
| Titanium nitride - TiCl4 - Purity based upon metals ( ppbv total) | tbd | tbd | tbd | tbd | tbd
Metal Gate Electrodes - Ru/RuO2
Metal Gate Electrodes - Ru/RuO2 - RuO4 [20427-56-9] (assay)
| CVD/ALD Thin Film Process Fluids - MIM capacitor Electrodes | tbd | tbd | tbd | tbd | tbd
MIM capacitor Electrodes - Ru / RuO2
MIM capacitor Electrodes - Ru(EtCp)2 [32992-96-4] (assay)
MIM capacitor Electrodes- TiN
| MIM capacitor Electrodes - TiN - TiCl4 [7550-45-0] (assay) | tbd | tbd | tbd | tbd | tbd
| MIM capacitor Electrodes - TiCl4 (Metals Purity ( ppbv total)) | tbd | tbd | tbd | tbd | tbd
| MIM capacitor Electrodes - TaN | >99.7 | >99.7 | >99.7 | >99.7 | >99.7
MIM capacitor Electrodes - TaN - PDMAT [19824-50-0]
| MIM capacitor Electrodes - TaN - TAIMATA [629654-53-1] (assay) | tbd | tbd | tbd | tbd | tbd
| MIM capacitor Electrodes - TaN - TBTEMT [ 511292-99-2] (assay) | tbd | tbd | tbd | tbd | tbd
| MIM capacitor Electrodes - WN | <100 | <100 | <100 | <100 | <100
| MIM capacitor Electrodes - WN - WF6 [7783-82-6] (assay) | >99.7 | >99.7 | >99.7 | >99.7 | >99.7
| CVD/ALD Thin Film Process Fluids - Interconnect Conductive Films | tbd | tbd | tbd | tbd | tbd
Conductive barriers - strike layers - Etch Stops - Caps - etc. (cmos and tsv's)
Conductive barriers - TaN - Tantalum Amides (e.g., PDMAT [19824-50-0], TBTDET [169896-41-7]; TAIMATA; TBTEMT)
| Tantalum Amides - (Critical specified metals/total metals (ppbw) ) | tbd | tbd | tbd | tbd | tbd
| Tantalum Amides (O2, H2O (ppbv)) | tbd | tbd | tbd | tbd | tbd
| Tantalum Amides (Hydrocarbons (ppm)) | tbd | tbd | tbd | tbd | tbd
Conductive barriers - Titanium Nitride Precursors - halides and Amides ; TiCl4; TDMAT; TEMAT; TDEAT; TiCp(X)3
| Titanium Nitride Precursors (Critical specified metals/total metals (ppbw) ) | tbd | tbd | tbd | tbd | tbd
| Titanium Nitride Precursors (O2, H2O (ppbv)) | tbd | tbd | tbd | tbd | tbd
| Titanium Nitride Precursors Hydrocarbons (ppm)) | <1000 | <1000 | <1000 | <1000 | <1000
Titanium Nitride Precursors -
| TDMAT [3275-24-9] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
| TDMAT - (Purity (based on metals analyzed ( ppmv)) | tbd | tbd | tbd | tbd | tbd
| TDEAT [4419-47-0] ( (assay % by NMR)) | <1000 | <1000 | <1000 | <1000 | <1000
| TiCl4 [7550-45-0] (assay) | tbd | tbd | tbd | tbd | tbd
| TiCl4 Purity based upon metals ( ppbv total) | tbd | tbd | tbd | tbd | tbd
| TEMAT [308103-54-0] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
Tantalum Nitride Precursors - Tantalum Amides (e.g., PDMAT [19824-50-0], TBTDET [169896-41-7]; TAIMATA; TBTEMT)
| Tantalum Nitride Precursors - (Nb (ppm)) | tbd | tbd | tbd | tbd | tbd
| Tantalum Nitride Precursors - (Metals each element (ppb) ) | tbd | tbd | tbd | tbd | tbd
| Tantalum Nitride Precursors - (Hydrocarbons (ppm)) | <100 | <100 | <100 | <100 | <100 | something shifted
| Tantalum Nitride Precursors - (Chloride (ppm)) | >99.7 | >99.7 | >99.7 | >99.7 | >99.7
| Tantalum Nitride Precursors - (Hydroloysis product (alkylamine)) | tbd | tbd | tbd | tbd | tbd
| PDMAT [19824-50-0] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
| TBTDET [169896-41-7] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
| TAIMATA [629654-53-1] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
| TBTEMT [ 511292-99-2] (assay % by NMR) | tbd | tbd | tbd | tbd | tbd
Conductive barriers - Ruthenium metal/ RuO2 - Ru carbonyls; Ru amidinates e.g., Ru(EtCp)2 [47, 48] [32992-96-4]; RuO4 [20427-56-9]
| Ru(EtCp)2 [32992-96-4] (assay) | tbd | tbd | tbd | tbd | tbd
| RuO4 [20427-56-9] (assay) | tbd | tbd | tbd | tbd | tbd
| Bulk Metal Interconnects - vertical - horizontal wiring - tsv's | tbd | tbd | tbd | tbd | tbd
| Tungsten metal - Tungsten Hexafluoride (WF6) [7783-82-6] (assay) | 99.9 | 99.9 | tbd | tbd | tbd
| Tungsten Hexafluoride (Critical specified metals/total metals (ppbw)) [19] | <1/1000 | <1/1000 | tbd | tbd | tbd
| Tungsten Hexafluoride (O2 (ppbv)) | <10 | <10 | tbd | tbd | tbd | usually no H2O spec
| Copper metal - Cu(hfac)(tmvs) [139566-53-3] ( assay) | tbd | tbd | tbd | tbd | tbd
CVD/ALD Thin Film Process Fluids - Semi-conducting Thin Films
Epitaxial Si and Ge - incuding SiGe alloys
| Epitaxial Si and Ge - Disilane [1590-87-0] (assay) | 99.999 | 99.999 | tbd | tbd | tbd
| High Temperature Si Epitaxy - SiHCl3 [59] [10025-78-2] (assay) | 99.95 | 99.95 | tbd | tbd | tbd
| High Temperature Si Epitaxy - SiHCl3 (O2 (ppb)) | tbd | tbd | tbd | tbd | tbd | not typically specified in TCS
| High Temperature Si Epitaxy - SiHCl3 (H2O (ppb)) | tbd | tbd | tbd | tbd | tbd | not typically specified in TCS
| High Temperature Si Epitaxy - SiHCl3 (Hydrocarbons (ppb)) | tbd | tbd | tbd | tbd | tbd | not typically specified in TCS
| High Temperature Si Epitaxy - SiHCl3 (Inboard He Leak Rate (atm cc/sec)) | tbd | tbd | tbd | tbd | tbd | not typically specified in TCS
| High Temperature Si Epitaxy - SiHCl3 (Dopants B, Al, P, As (ppb)) | <0,5 | <0,5 | tbd | tbd | tbd | new
| High Temperature Si Epitaxy - SiHCl3 (Metals like Fe (ppb)) | <1 | <1 | tbd | tbd | tbd | new
| High Temperature Si Epitaxy - SiHCl3 (Carbon (ppb)) | <500 | <500 | tbd | tbd | tbd | new
| High Temperature Si Epitaxy - SiH4 [7803-62-5] (assay) | tbd | tbd | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 [7782-65-2] (assay) | tbd | tbd | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Carbon Dioxide ( ppmv)) | <100 | <100 | <100 | <100 | <100
| Epitaxial Si and Ge - GeH4 (Carbon Monoxide ( ppmv)) | <100 | <100 | <100 | <100 | <100
| Epitaxial Si and Ge - GeH4 (Digermane ( ppmv)) | <100 | <100 | <100 | <100 | <100
| Epitaxial Si and Ge - GeH4 (Nitrogen ( ppmv)) | <10E-9 | <10E-9 | <10E-9 | <10E-9 | <10E-9
| Epitaxial Si and Ge - GeH4 (Methane(ppmv)) | tbd | tbd | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Oxygen + Argon (ppmv)) | tbd | tbd | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Trigermane( ppmv)) | <2 | <2 | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Water( ppmv)) | <1 | <1 | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Chlorogermanes( ppmv)) | <20 | <20 | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Germoxanes( ppmv)) | <2 | <2 | tbd | tbd | tbd
| Epitaxial Si and Ge - GeH4 (Hydrogen(ppmv)) | <1 | <1 | tbd | tbd | tbd
| CMOS Channel Materials | [5 empty] | ? Does this fit here
| Epitaxial Si Channel Replacement films | <0.5 | <0.5 | tbd | tbd | tbd
| Epitaxial Si Channel Replacement - GaAs / InGaAs | <1 | <1 | tbd | tbd | tbd
| GaAs / InGaAs - Ga(CH3)3 (TMG) [1445-79-0] (assay) | <1 | <1 | tbd | tbd | tbd
| Ga(CH3)3 (TMG) (Metals each element( total ppb)) | <5 | <5 | tbd | tbd | tbd
| GaAs / InGaAs - In(CH3)3 (TMI) [3385-78-2] (assay) | <50 | <50 | tbd | tbd | tbd
| GaAs / InGaAs - arsine ( AsH3) [7784-42-1] (assay) | tbd | tbd | tbd | tbd | tbd
Epitaxial Si Channel Replacement - GaN
| Epitaxial Si Channel Replacement - GaN - TMG [1445-79-0] (assay) | 99.999 | 99.999 | tbd | tbd | tbd | values are from GaN Epi for power SC!
| Epitaxial Si Channel Replacement - GaN - NH3 [7664-41-7] (assay) | 99.9999 | 99.9999 | tbd | tbd | tbd
| GaN - NH3 (Iron (ppbw)) | <100 | <100 | tbd | tbd | tbd
| Epitaxial Si Channel Replacement - SiC | tbd | tbd | tbd | tbd | tbd
| Channel Replacement - SiC - SiH4 [7803-62-5] ( see above) | 99.9999 | 99.9999 | tbd | tbd | tbd
| Channel Replacement - SiC - C3H8 (propane) [74-98-6] (assay) | 99.999 | 99.999 | tbd | tbd | tbd
| SiC - C3H8 (propane) - (Total metals (ppbw)) | tbd | tbd | tbd | tbd | tbd | need to verify - not sure it is relevant
Channel Stress-Strain Management Films
Strain Management Films - SiGe alloys
Low Temperature Si/Ge Epitaxy
Low Temperature Si/Ge Epitaxy - Dichlorosilane (DCS), diluteB2H6, dilute PH3, dilute monomethylsilane
| Low Temperature Si/Ge Epitaxy (O2 (ppb)) | <500 | <500 | tbd | tbd | tbd
| Low Temperature Si/Ge Epitaxy (H2O (ppb)) | <200 | <200 | tbd | tbd | tbd
| Low Temperature Si/Ge Epitaxy (Hydrocarbons (ppb)) | <1000 | <1000 | tbd | tbd | tbd
| Low Temperature Si/Ge Epitaxy - Inboard He Leak Rate (atm cc/sec) | tbd | tbd | tbd | tbd | tbd
Point of Use Purification
POUF - SiH4, Si2H6, GeH4, process and purge H2, process and purge N2
| POUF - SiH4, Si2H6, GeH4 (O2 (ppb)) | <10 | <10 | <10 | <10 | <10
| POUF - SiH4, Si2H6, GeH4 (H2O (ppb)) | <20 | <20 | <20 | <20 | <20
| POUF - SiH4, Si2H6, GeH4 (Hydrocarbons (ppb)) | <20 | <20 | <20 | <20 | <20
| POUF - SiH4, Si2H6, GeH4 (Inboard He Leak Rate (atm cc/sec)) | <10E-10 | <10E-10 | <10E-10 | <10E-10 | <10E-10
| POUF - Deposition gases - GeH4 [7782-65-2] | No | No | No | No | No
| POUF - GeH4 (O2, H2O (ppbv)) | <10 | <10 | <10 | <10 | <10
| POUF - GeH4 (Other dopants (ppbv)) | <20 | <20 | <20 | <20 | <20
| Mixing tolerance for mixtures | <20 | <20 | <20 | <20 | <20
| GeCl4 [10038-98-9] | <10E-10 | <10E-10 | <10E-10 | <10E-10 | <10E-10
| GeH4 [7782-65-2] | Yes | Yes | Yes | Yes | Yes
| SiH2Cl2 [4109-96-0] | tbd | tbd | tbd | tbd | tbd
| Si2H6 [1590-87-0] | < 500 | < 500 | < 500 | < 500 | < 500
| VIc2 Si3N4 | <1 | <1 | <1 | <1 | <1
| SiH4 [7803-62-5] | +/-1% | +/-1% | +/-1% | +/-1% | +/-1%
| Si2H6 [1590-87-0] | tbd | tbd | tbd | tbd | tbd
| Other available High purity volatile liquid precursors available for Electronic Thin film applications (General) | tbd | tbd | tbd | tbd | tbd
| Deposition Gases CVD/PECVD/ ALD (General) | tbd | tbd | tbd | tbd | tbd
| Deposition Gases CVD/PECVD/ ALD (General) (O2, H2O (ppbv)) | tbd | tbd | tbd | tbd | tbd
| Deposition Gases CVD/PECVD/ ALD (General) (Critical specified metals/total metals (ppbw) )[19] | tbd | tbd | tbd | tbd | tbd
| Deposition Gases CVD/PECVD/ ALD (General) (Dopants (ppbv)) | tbd | tbd | tbd | tbd | tbd
| Metal Dopants (e.g., FeCp2, MgCp2 ) (assay) | tbd | tbd | tbd | tbd | tbd
| Amorphous Carbon (Double Patterning) e.g., propylene [9003-07-0] and acetylene [74-86-2] | tbd | tbd | tbd | tbd | tbd
| Amorphous Carbon (Double Patterning) (Other hydrocarbons (ppmv)) | tbd | tbd | tbd | tbd | tbd
| Amorphous Carbon (Double Patterning) (Total metals (ppbw)) | < 1000 | < 1000 | < 1000 | < 1000 | < 1000
Deposition gases (e.g. NH3)
| Deposition gases (e.g. NH3) (O2, H2O (ppbv)) | <500 | <500 | tbd | tbd | tbd
| Deposition gases (e.g. NH3) (Critical specified metals/total metals (ppbw)) [19] | <100 | <100 | tbd | tbd | tbd
| Deposition gases (e.g., N2O, NO) | tbd | tbd | tbd | tbd | tbd
| Deposition gases (e.g. NH3) (O2, H2O (ppbv)) | <100 | <100 | tbd | tbd | tbd
| Deposition gases (e.g. NH3) (Critical specified metals/total metals (ppbw)) [19] | <200 | <200 | tbd | tbd | tbd
Deposition gases - electrical dopants (e.g. AsH3, PH3, B2H6)
| electrical dopants (O2, H2 (ppb)) [36] | <100 | <100 | tbd | tbd | tbd
| Other dopants (ppbv) | <1 | <1 | tbd | tbd | tbd
| Other dopants (Mixing tolerance for mixtures (relative variance)) | +/-1% | +/-1% | tbd | tbd | tbd
Implant gases - AsH3, PH3, BF3
| Implant gases (O2, H2O (ppbv)) | <500 | <500 | <500 | <500 | <500
| Implant gases (Other dopants (ppbv)) | <1 | <1 | <1 | <1 | <1
| Implant gases (Mixing tolerance for F2 (relative variance)) | +/-4% | +/-4% | +/-4% | +/-4% | +/-4%
| Implant gases (Other constituents (ppbv)) | <25000 | <25000 | <25000 | <25000 | <25000
Laser gases - Litho, (e.g. F2/Kr/Ne)
| Laser Gases (O2, H2O (ppbv)) | <500 | <500 | tbd | tbd | tbd
| Critcal Components [46][48] | [5 empty] | Focus is for HP polymer materials used in liquid delivery systems | Update SEMI F57
HP materials used in distribution of UPW
| Total organic carbon (ng/m2) [47] | 60000 | 30000 | 30000 | 30000 | 30000 | baseline current performance of HP polymer materials
| Critical Organics as C (ng/m2) | tbd | tbd | tbd | tbd | tbd
| Number of particles >critical particle size (#/l or #/actuated component) [45], | 1000 | 1000 | 1000 | 1000 | 1000 | Non EAP. For additional information refer to UPW notes. | Update SEMI F104
| Metals (ug/m2) (Al, As, Ba, Cd, Co, Cr, Cu, Fe, K, Li, Mg, Mn, Mo, Na, Ni, Pb, Sb, Sn, Sr, Ti, V, W, Zn) | <10 | <10 | <10 | <10 | <10 | baseline current performance of HP polymer materials
| Metals (ug/m2) Ca | <20 | <20 | <20 | <20 | <20 | baseline current performance of HP polymer materials
| Anions (ug/m2) Cl , NO2, NO3, PO4, SO4 | 300 | 300 | 300 | 300 | 300 | baseline current performance of HP polymer materials
| F | 60000 | 10000 | 10000 | 10000 | 10000 | baseline current performance of HP polymer materials
HP materials used in distribution of Chemicals
| All cleaning chemistries (aqueous and solvent): number of particles /l > critical particle size [45] | 1000 | 1000 | 1000 | 1000 | 1000 | Using 10% of the value for chemical specification for non EAP
| All cleaning chemistries (aqueous and solvent) Al, Ba, B, K, Na, K, Fe, Li, Mg, Mn, Pb, Sn, Ti, Zn, Ni, Cu, Cr, Co, (ug/m2, each) | <10 | <10 | <10 | <10 | <10 | baseline current performance of HP polymer materials
| Acid and Caustic Delivery Systems Ca | <20 | <20 | <20 | <20 | <20 | baseline current performance of HP polymer materials
| 100% IPA: Al, Ba, B, K, Na, K, Fe, Li, Mg, Mn, Pb, Sn, Ti, Zn, Ni, Cu, Cr, Co, Ca (ug/m2, each) | <10 | <10 | <10 | <10 | <10 | baseline current performance of HP polymer materials
| 100% IPA Ca | <20 | <20 | <20 | <20 | <20 | baseline current performance of HP polymer materials
| 100% IPA: High molecular weight organics (ug/m2) | tbd | tbd | tbd | tbd | tbd
| 100% IPA: Anions (ug/m2) | 300 | 300 | 300 | 300 | 300 | baseline current performance of HP polymer materials
| Acid and Casutic Delivery metals ng/m2 (Al, As, Ba, Ca, Cd, Co, Cr, Cu, Fe, K, Li, Mg, Mn, Mo, Na, Ni, Pb, Sb, Sn, Sr, Ti, V, W, Zn) | <15 | <15 | <15 | <15 | <15 | baseline current performance of HP polymer materials
| Post-CMP clean chemicals: particles >0.2 µm (/m2) Surfactant Based | 10 | 10 | 10 | 10 | 10 | baseline current performance of HP polymer materials

## Notes for Table YE3
INDEX
Notes for Table YE3
[1] Logic was chosen to be a driver for Yield roadmap for the following reasons: • It has tighter pitch size • Flash is moving to 3D structure where the pitch equivalent is much larger • Logic has many more metal layers and higher complexity, increasing risks • Logic has less redundancy compared to Flash memory • Flash has self-alignment – none for logic Critical particle size is based on ½ pitch design rule for electrically active particles, assuming that "bridging" can be caused by particles smaller than the ½ pitch size. All other particles are not expected to become a device "killer" until their size exceeds ½ pitch.
[2] Airborne particle requirements are based on ISO 14644-1 at “at rest”.[3]
[3] Ion/species indicated is basis for calculation. Exposure time is 60 minutes with starting surface concentration of zero. Basis for lithography projections is defined by lithography tool suppliers. Metals and organics scale as defined in the surface preparation roadmap for metallics and organics. Values listed in table are based on experience, however, all airborne molecular contaminants can be calculated as S=E*(N*V/4); where S is the arrival rate (molecules/second/cm2), E is the sticking coefficient (between 0 and 1), N is the concentration in air (molecules/cm3); and V is the average thermal velocity (cm/second). The following sticking coefficients have been proposed; SO4 = 1x10 -5, NH3 = 1x10 -6, Cu = 2x10 -5. The sticking coefficients for organics vary greatly with molecular structure and are also dependent on surface termination.
[4] Includes P, B, As, Sb
[5] Contaminant targets apply up to POE (point-of-entry). POE is defined as the entry point to the equipment or subequipment, see also the text. Benchmark data has been collected both at Point of Delivery (POD) or Point of Entry (POE), which typically show only minor differences.
[6] Critical metals and ions may include: Al, As, Ba, Ca, Co, Cu, Cr, Fe, K, Li, Mg, Mn, Na, Ni, Pb, Sn, Ti, Zn. Three different case studies were reviewed where the levels of Ca, Fe, and Ni in the UPW resulted in levels of problem densities (atoms/sq cm) on the wafer. These were reduced to acceptable levels by reducing the level of these elements in the UPW to levels well below 10 ppt. In only one case does the data exist that showed success by obtaining values below 0.5 ppt. These results drive the 1.0 ~ 0.5 ppt values.
[7] Units on all contaminants in WECC Table are often given as ppb (or ppm or ppt, we use ppb here solely for demonstration purposes). The reader should be aware that these units of parts per billion (ppb) may be ppb by mass, volume, or molar ratios. Where not designated, the following guidelines apply: Chemicals and UPW are typically ppb by mass, while gases and clean room are typically ppb by volume. In the case of the fluid acting as an ideal gas, ppb by volume is equal to ppb molar. The notable exception to the above is metals in gases that are ppb by mass. In UPW, ppb is used based on mass of water. Since 1 liter of water equals ~1 kg, 1 ppb is equal to 1 µg/l. Some parameters in the tables may be considered process variables rather than contaminants in the classical meaning. They are marked by an asterix. The limits are sometimes fluent.
[8] Detection of metals at the levels indicated will be dependent on sampling time and flow rate. Sticking Coefficients vary widely for metals. It is generally believed that Cu has a sticking coefficient 10x of other metals, and therefore the guideline for Cu could be lower.
WAS | [9] Key particle size for scratching particles depends on mean particle size of slurry. Target level will be specific to slurry and wafer geometry sensitivity.
IS | [9] For Mullti-Modal abrasives, values are based on the MPS of the largest Mode. This is to measure the amount agglomeration that cause scratch defects during the CMP process.
[10] The Dissolved Nitrogen range is solely for the physical process needs of megasonics cleaning. Processes without megasonics cleaning can ignore the line item. The concentration is process specific and needs to be determined by the end user. Factors to consider include UPW temperature, partial pressure in the gas phase and megasonic energy input at the tool. Other gases, such as oxygen and hydrogen, may be used with different optimum levels. Process enhancements through chemistry associated with the other gases or other chemicals are outside of the scope of this chapter. While a relatively wide range of dissolved nitrogen values may be acceptable based on other process variables it is important to note that tight control at the desired value is warranted. DN2 spec is considered at POE, taking into account possibility that N2 control inside the tool.
[11] As of the current year's update the finest production available sensitivity liquid particle sensor for chemicals is 0.040 µm, although measurement technology is being developed at 0.020 µm. Values obtained by these particle counters are not directly comparable to the roadmap values and need to be normalized to critical particle size values in the roadmap using the equation and methods of Footnote A above. Interim solution to higher sensitivity particle counter is to collect data over longer time period to provide greater precision in the data near the threshold sensitivity of the counter. Most benchmark data has been collected at Point of Delivery (POD) or Point of Entry (POE) and is the basis for parameters. This particle counting efficiency is not taken into account, which may vary with particle counter model, particle materials, chemical, etc.. It is based on actual data reported.
[12] SMC Organics: Single wafer shall be oxidized to make organic-free, then wafer shall be exposed for 24 hours and top side analyzed by TD-GC-MS with 400°C thermal desorption, and quantitation based on hexadecane external standard. TIC response factor per SEMI MF 1982-1103 (formerly ASTM 1982-99).[4] Limits determined by above method are a guideline for many organics. Note higher limits can be used for process wafers oxidized or cleaned prior to subsequent process step. Processes such as gate oxide formation, or polysilicon deposition, may be more sensitive to organics, especially high boilers such as DOP. Silicon nitride nucleation may also be more sensitive than above for some processes. Please note dopants requirement is covered in earlier section. Contamination levels are time based, and samples should be exposed for a weeks time for better sensitivity; ng/cm2/week. Total contamination level on reticles that cause problems also vary with energy exposure. These guidelines subject to change with new data currently being generated.
[13] SMC Dopants: Single wafer is first stripped with HF to yield dopant-free surface and than exposed for 24 hours. Topside of wafer is analyzed by methods known to give reliable recovery of boron. This is a guideline for dopants based on sampling in operating running fabs. Lower specifications may be required for key FEPs, especially for smaller geometries, lower thermal budgets, and for lightly-doped devices. If wafers are stripped with HF or BOE immediately prior to next thermal process, then steps may become less sensitive to surface molecular dopants, and higher limits apply. Note that BEPs tend to be orders of magnitude less sensitive to dopants than FEPs.
[14] SMC Metals: Single wafer known to meet the ITRS FEP spec of 1E10 atoms/cm2, from the Starting Materials table, is exposed to a clean environment for 24 hours. Subsequent analysis of top surface by VPD-ICP-MS or VPD-GFAA. Lower specifications may be required for key FEPs, especially for smaller geometries. If wafers are cleaned prior to the next thermal process, then air exposure during earlier steps may be less of an issue. Note that majority of environmental metallic contaminants are particles, not molecular. If total particles on wafers are kept in spec than majority of metals, most metals from the environment should be within specifications. Back-end processes (BEPs) tend to be less sensitive to metals that FEPs provided not particles. Specs of twice the incoming wafer specs are readily achievable and readily measurable in case of wafers exposed for 24 hours.
[15] SMC General: A 24-hour exposure will accentuate the contamination per wafer as wafers are often exposed too much shorter times in actual processing. The above SMC (surface molecular contamination) limits are preliminary, and no single value applies to all process steps or types of organics, dopants or metals. The SMC limits can vary substantially from process to process, and local air purification or purges may be needed to control contaminant levels.
[16] Dissolved oxygen (DO) has an effect on pre-gate oxide cleaning and the etch rate of non H-passivated SiO2 and copper structures. The level in the table is that of the most stringent. It is expected that slightly higher levels within the same order of magnitude would not have any significant effect on manufacturing processes. If the water for a specific processes need to remain at low oxygen concentrations lower levels of dissolved oxygen could provide somewhat larger process time windows before critical concentration levels are reached. It is known that some fabs consider DO a process variable and operate at DO levels 3 orders of magnitude higher than stated in the table. Corrosion rates as a function of DO are not a linear relationship for all materials, specifically copper etch rates are near a maximum at 300 ppb DO. DO spec is considered at POE, taking into account possibility that N2 blanketing may be provided at the tool level.
[17] Uncertain at this time what target levels might be set given the variety of chemistries used in the industry and unknown sensitivity of the wafer to particles or ionic contamination in the chemical. This parameter is identified as a potentially critical one that should be considered and work is ongoing to define the correct levels.
[18] The value of silica used in the roadmap is at current low practical detection limit and it is supposed to indicate the overal health of the system. Silica may occur in UPW in a form of reactive (dissolved) or colloidal silica. Colloidal silica is considered detrimental for the semiconductor manufacturing. However, the killer particle size of currently considered generations became so small, making it impracticle to expect that the colloidal silica might be measured as delta between total and reactive silica in UPW. Hence, colloidal silica is considered to be controlled as particles down to non-detectable (as silica) level. As a result only reacive silica may occur in UPW. For practicle purposes either of the tests total or reactive silica can be used. However, measuring both total and reactive silica may not be considered to confirm presence of colloidal silica. Inhigh concentration (1 ppb), dissolved Silica is a potential cause for water spotting. When boron is controlled below or near limits of boron detection the level of dissolved silica is usually bellow the limits of detection as well. The silica specification is considered at POC (point of connection) based on the industry experience – no water marks were reported at the specified level. It is expected that the wafer will leach silica, adding silica level on the top of UPW level.
[19] The list of critical metals (e.g., Al, Ca, Cu, Fe, Mg, Ni, K, Si, Na) varies from process to process depending on the impact on electrical parameters such as gate oxide integrity or minority carrier lifetime as well as mobility of the metal in the substrate. The metals listed in note [G] for liquid process chemicals are of concern but the issues around metals in specialty gases are primarily around the potential for corrosion to add metal particles to the gas flow (e.g., Fe, Ni Co, P). The potential for volatile species containing metals must be considered for each specialty gas but are generally not present in the bulk gases.
[20] The following is a complete list of metal ions of concern in certain liquid chemicals: Ag, Al, As, Au, Ba, Ca, Cd, Co, Cr, Cu, Fe, K, Li, Mg, Mn, Mo, Na, Ni, P, Pb, Pd, Pt, Ru, Sb, Sn, Sr, Ti, V, W, Zn.
[21] Elements listed that are not in parentheses may cause high or some risk to device quality and may often be present in process chemicals. Elements listed that are in parentheses may cause high risk to device quality but are not typically present in process chemicals.
[22] TOC on-line measurement remains the reference detection method for the presence of total organics in UPW. Users of this roadmap are encouraged to speciate organics in their UPW. For the practical purposes, the on-line TOC measurement may be considered as critical organics, unless non-critical fraction is determined. Less than 1 ppb of Total Organic Carbon was proven successful for immersion lithography operation, concerned about lens hazing. Total Organic Carbon in UPW is defined as equivalent amount of CO2 produced from UV oxidation at 185 nm wavelength inside of the UV reactor of a TOC monitoring device. Currently used online, TOC analyzers are non-specific in their measurement of organic contaminants. Some organics are expected to have higher impact to the semiconductor manufacturing than others, hence TOC specification has been replaced with critical and non-critical organics for all applications, except for immersion photolitography. Critical Organics are those with the highest boiling point. Organics with high boiling points could deposit on the wafer surface in a high temperature furnace and result in the formation of silicon carbide. This is considered a fatal defect at levels of 1.0E+12 atoms per cm2 as C. The boiling point of organic compounds is used to categorize them as a “critical”. At this time, any organic compound with a boiling point higher than 200°C is considered critical. Based on past experience acceptable concentration of total critical organic compounds is <1.0 ppb (as measured by on-line instrumentation). Non critical organics is proposed to be as high as 3.0 ppb. Modeling and experimental analysis are in progress to validate correlations of critical organic compounds to device defects. The end user may want to consider POU solutions for immersion litho if 1.0 ppb TOC water quality is chosen to be controlled for the immersion lithography tools, independently of the type of the organics comprising TOC. In addition to the boiling point characteristics, size of the organic molecule becomes another important consideration. Very high molecular weight dissolved organics may form particles on the wafer as the water dries out. This phenomenon was confirmed experiementally refer to supplemental information) and hence the particle related concern is growing as dimensions become smaller. At this point 500,000 dalton is considered to be the size of the organic molecule potentially forming killer size particles. The color in the ITRS table is provided to indicate the need of metrology development for online measurement of critical organics, as well additional work needed to better define the critical organics via FMEA. The organics specification is defined at the POE (Point of Entry) to the manufacturing tools, assuming that control of the organics level is provided outside of the tools, whereas the manufacturing process technology takes into account contribution by the tool materials and chemistry used.
[23] The photolithography AMC guidelines are for tools with ArF lasers only, and are based on inputs from the Litho TWG. The critical metrology guidelines are from the metrology TWG. All photolithography tools and some metrology equipments have chemical filters on the makeup air to the internals of the tools. These filters have a finite lifetime, which is dependent on the contaminant loading. Providing a chemically cleaner environment will extend the life of these filters.
[24] Other critical ions may include inorganic ions such as Fluoride, Chloride, Nitrate, Nitrite, and Bromide, Sulfate as well as ammonium. However no reference was currently found that these ions in typical concentrations found in ultrapure water up to 50 ppt have any impact on the process. Phosphate specification level was defined at 20 ppt based on higher risk of deposition. Also for organic anions such as acetate, formate, propionate, citrate, and oxalate no harmful levels have been established up to now. The ions specification is considered at POP (point of process) due to assumption that both UPW system as well as POU filters may be used to control metals, while tool environment may contribute metal contamination. Ammonium, sulfate, and nitrite can be contributed by air.
[25] The variety of CVD and ALD precursors is continuously increasing as well as their applications. The contaminant types and levels vary widely due to the different chemical behavior. An overview about typical precursors is therefore given in attachment Precursor table.
[26.1] The maximum allowed particle concentration in UPW has been updated based on the following assumptions and considerations: A. The maximum concentration of particles in UPW is a function of the allowed defect level on the wafer as defined in the More Moore table and particle deposition form UPW to the wafer surface. The reference point will be the critical particle size. The relationship between acceptable particle concentrations on the wafer and that in UPW is based on experimental data (see supplemental materials for reference) and theoretical modeling. The user of this part of the table is encouraged to refer to the supplemental materials for specific details of the calculations. B. Significant metrology hurdles for particles at the critical size still exist. Given the fact that the metrology gap is growing and has not been addressed for last decade, the risk management for particle control is being shifted towards the effective mitigation measures, including particle filtration and reduction of the particle challenge to the final filters. This idea is being incorporated with a SEMI task force produced guides ([5] SEMI C079) for filter performance validation, as a method to control the particles on the wafer surface and mitigate the risk of particle contamination impacting yield. A similar task force has developed a test method for evaluation of the cleanliness of the virgin ion exchange resin, SEMI C093. [6]. [26.2] Colloidal silica was previously added to emphasize the criticality of colloidal silica particles in UPW, than the currently measured reactive and total silica (Si) in water. Existing UPW metrology is not sensitive enough to allow the detection of silica particles at the specified size as a chemical. Silica particles are believed to have the ability to adsorb metals and act as a transport mechanism to potentially deposit metals on the surface of a wafer. Such metal deposition could seriously impact yield. There is no metrology that can effectively detect colloidal silica directly. And therefore, the colloidal silica specification was merged with the particles. [26.3] An additional particle requirement has been specified for EUV mask production, which requires even tighter particle control. However the allowed concentration is an order of magnitude lower than that specified for high volume manufacturing facilities (refer to supplemental information for additional details). The killer size of the particles for this application is a function of the EUV wavelenght. [26.4] Supplemental information shows different mechanisms of the effect of the particles to the device. Bridging the current is considered to be more critical one as even smaller than other EAP may kill the device. The particles specification is considered at POP (point of process) due to assumption that both UPW system, critical components, as well as tool environment may generate particles, taking into account capabilities of both UPW delivery system (including POU filters) and the tool environment contribution. However, for the practical purposes, the assumption​ is that most particles will be contributed by the tool, thus requiring even higher UPW quality targets at POE.
[27] It needs to be considered that the total H2O2 anion concentration will impact the life time of the solution. Also the fluoride in the ppm range of the total chemical mixture can etch the wafer.
[28] Concentrations higher than 100 ppb could cause corrosion especially in back end of line processes.
[29] The ultrapure water parameters provided in this table are applicable for the most critical process unless otherwise identified by additional footnotes. Further information can be found in the Supplementary tables.
[30] Typical Organic Acids found in cleanroom environments that may be of concern include Acetate, Citrate, Formate, Glycolate, Lactate, Oxalate, and Proprionate. Others may also be of concern. These acids can be a significant load on acid removal filters.
[31] Ideally, continuous monitoring using online instrumentation would be preferred when practical since this can give both long term averages and catch excursions. When online monitoring is not available, an average grab sample for at least 4 hours, and not more than 24 hours is recommended, to get an average, increase sensitivity of the analysis, and avoid short term transient effects.
[32] Other corrosive species include contaminants such as chlorine. Humidity is also of major concern, as it exacerbates corrosion. The humidity should be kept as low as possible in corrosive environments.
[33] Calculations for expressing ng/L into ppt are; [(ng/(L of Air) * (24.4 L of Air)/mol Air / MW(ng/nmol)) * *1000picomol/nmol ] = picomol/mol of Air = ppt molar and/or ppt volume.
[34] For certain processes such as sputtering POE purifiers may be required for N2 and Ar
[35] CO2 here is assumed to be used for wet cleaning and other equipment, not for super critical CO2 applications or dry etching.
[36] Epi – need to purify @ 45 nm; currently must add purifiers from B2H6/Germane/PH3/AsH3 – need 100 ppb
[37] The variation is defined at one location over time in at rest conditions. As reference point for the POE a location is chosen 0.3 m below the ceiling panels. Common sense requires that sensitive equipment are not installed heat sources within the cleanroom, since they may impact the temperature control between the reference point and the actual inlet to the mini-environment/tool filter.
[38] This temperature stability requirement is for immersion photolithography tools, using UPW as an immersion fluid, and based upon utility requirements projected by some tool manufacturers in 2005. It represents the maximum rate of change and stability of the temperature of the cold UPW supplied to the tool in order for the tool to maintain process required temperature stability.
[39] The list of metals was increased to 24 elements to include additional elements listed in 2013 FEP table. Values based on FEP definitions of criticality to process. Users of the Table are strongly encouraged to review the Metal White Paper in supplemental materials for additional information. Metals are considered to be particularly important for CMOS sensors – special parameter was added to the roadmap. The metal specification is considered at POP (point of process) due to assumption that both UPW system as well as POU filters may be used to control metals, while tool environment may contribute metal contamination.
[40] Value of 100 pptv refractory compounds is not to be seen as an absolute limit but a recommanded value derived from total dose considerations. Based on experience bottom lens deterioration from refractories is prevented for 7–10 years if an annual integral dose of 8.8E+5 ppth of refractories is not exceeded. This calculates back to an average permanent sum concentration of 100 pptv. Actual concentrations can vary in a bandwidth of factor 3-5 around this value.
[41] The resin by-products are only a small part of the organics in H2O2, but have shown the most serious effects on the wafer. The organics leached out from the resins are different depending on the type of resin, but have shown a much higher sticking potential than the other solvents in the H2O2. Total amines are mentioned here as an example only. Measurement of these organics is possible, e.g. by ion chromatography or liquid chromatography with mass spectrometer, but there a variety of problems, such as the sample preparation since these compounds can be very volatile/sensitive.
[42] Hydrogen peroxide (H2O2) was confirmed to be produced by UV (ultraviolet) lamps in the polish loop at the levels of up to 30 ppb, the level exceeding dissolved oxygen, which is also considered to be an oxidant. The effect of specific level of H2O2 to the wafer has not been fully validated yet. Experimental work conducted by Screen (see supplemental information) indicated that 10 ppb did not result in higher corrosionrate of cobalt, which is considered to be most susceptible to corrosion. Given that DO2 limit is set at 10 ppb, it has been decided to align between DO2 and stronger oxidant H2O2. Further research is needed to confirm UPW spec levels. It is possible that no significant difference in corrosion rate between <1 ppb H2O2 and 10 ppb of that used on the experiement was due to the fact that cobal was already affected by O2 in the environmnet even at lower lever of H2O2. H2O2 control is considered at POD (point of distribution) based on the assumption that no H2O2 production is expected downstream that point.
[43] More than one IDM regards the combination of 0.1 ppb limit value and the refractory compounds definition as not supported by own experience on lens integrity; General opinion is that light boiling perfluorated compounds (normally from chiller thermofluids) (e.g. Galden LT) do not have lens effects due to their lens sticking probability; A severeness rating is requested stating that Silanol and Siloxane compounds are of high importance in their contribution, S- and P- compounds following and of least importance perhalogenated compounds
[44] There is an opinion that the limit value of 0.5 ppbv total inorganic acids is too strict; Experience exists that 2 ppbv would be an acceptable limit value maintaining good yield/no or low defect conditions
[45] For particles for critical components it is #/l in UPW for tubing and non actuated components (like filters) or it is per actuated component (like a valve) using SEMI F104. This is 10% of UPW specification.
[46] Critical components includes all high purity components in the flow stream for liquide delivery and is focused on the last filter and beyond (i.e., close to wafer). This includes pipe, tube, valves, control elements, filters that are intergral to final delivery.
[47] Based on semi F40/57 test. Values based on current perfromance of HP materials.
[48] Effect of any contaminants in ciritical components on wafer is not yet understood.
[4] SEMI MF1982-1103 (previously ASTMF 1982-99e1), Standard Test Methods for Analyzing Organic Contaminants on Silicon Wafer Surfaces by Thermal Desorption Gas Chromatography, SEMI.
[5] SEMI C79-0113 GUIDE TO EVALUATE THE EFFICACY OF SUB-15 nm FILTERS USED IN ULTRAPURE WATER (UPW) DISTRIBUTION SYSTEMS. Www.semi.org
[6] SEMI C093: GUIDE FOR DETERMINING THE QUALITY OF ION EXCHANGE RESIN USED IN POLISH APPLICATIONS OF ULTRAPURE WATER SYSTEM
NEW | [49] Chemistry containing Micelles will require a larger particle size threshhold to be used to prevent counting Micelles as harmfull particles.
NEW | [50] Chemistry containing Micelles will require a larger particle size threshhold to be used to prevent counting Micelles as harmfull particles.

## Table YE4
INDEX
Table YE 4 AMC Monitoring Methods
| Year of Production | 2017
| Flash ½ Pitch (nm) (un-contacted Poly)(f) | 13
| DRAM ½ Pitch (nm) (contacted) | 20
| MPU/ASIC Metal 1 (M1) ½ Pitch (nm) | 25.3
| MPU Printed Gate Length (nm) †† | 17.6531
| MPU Physical Gate Length (nm) | 14.0028
Wafer Environment Control such as Cleanroom, SMIF POD, FOUP, etc….not necessarily the cleanroom itself but wafer environment.
| Critical particle size (nm) [1] | 10
| Number of particles (/m3) [1] [2] | ISO CL1 | Measurement methods
| Airborne Molecular Contaminants in Gas Phase (pptV, V for Volume)) [3, 7, 12,13,14,15,33]. |  | Online Monitoring |  | Wet or Dry Impinger Sampling [11] |  |  |  |  | Passive/Diffusive Sampling (Beaker/Dry trap) [12] |  |  |  | TD GC-MS |  |  |  |  | Surface analysis
| Lithography POI (wafer stage & reticle library) | Limits from YE3 table | Method | LDL | Method | LDL (pptV) | Sample volume (L) | Sample flow rate (LPM) | Sample frequency | Method | LDL | Sample duration (d) | Sample frequency | Method challenges | LDL (pptV) | Volume (L) | Sample flow rate (mL/min) | Sample frequency | Method | LDL
| Total Inorganic Acids | 100 | PWPD-IC [2] *) | 10 | IC-Conductivity [13,14] | 10-100 | 1500-200 | 1-5 | Every 2 months | IC-Conductivity [13,14,16] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
| Total Organic Acids [30] | 100 | PWPD-IC [2] *) | 30 | IC-Conductivity [13,14] | 10-100 | 1500-200 | 1-5 | Every 2 months | IC-Conductivity [13,14,16] | 100 | 7-90 | quarterly | Use special GC column | 100 | 20 to 40 | 50 to 200 | Every 2 months
|  |  | PTR-MS [5] *) | 200
| Total Bases | 200 | PWPD-IC [2] *) | 10 | IC-Conductivity [13,14, 15] | 10-100 | 1500-200 | 1-5 | Every 2 months | IC-Conductivity [13,14,15,16] | 100 | 7-90 | quarterly | For organic amines only (not total bases). Cut at C28. | 100 | 20 to 40 | 50 to 200 | t.b.d
|  |  | PAS [6] *) | 20
|  |  | IMS [1] | 1000
|  |  | CRDS [3] *) | 4000
| Refractory compounds (organics containing for example S, P, Si) | 50 | PTR-MS [5] *) | 50 | [9 empty] | Breakthrough to be studied for low boiling point S or Si compounds. Cut at C28. | 100 | 20 to 40 | 50 to 200 | t.b.d
Lithography POI (inspection tools stage)
| Total Inorganic Acids | 100 | PWPD-IC [2] *) | 10 | IC-Conductivity [13,14] | 100 | 200-500 | 1-5 | Every 6 months | IC-Conductivity [13,14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
| Total Organic Acids [30] | 100 | PWPD-IC [2] *) | 30 | IC-Conductivity [13,14] | 100 | 200-500 | 1-5 | Every 6 months | IC-Conductivity [13,14] | 100 | 7-90 | quarterly | Use special GC column | 100 | 20 to 40 | 50 to 200 | 2x annual
|  |  | PTR-MS [5] *) | 200
| Total Bases | 200 | PWPD-IC [2] *) | 10 | IC-Conductivity [13,14, 15] | 100 | 200-500 | 1-5 | Every 6 months | IC-Conductivity [13,14, 15] | 100 | 7-90 | quarterly
|  |  | PAS [6] *) | 20
|  |  | IMS [1] | 1000
| Condensable organics (definition to SEMI F21-1102, bp 150 °C) | 100 | PTR-MS [5] *) | 50 | [9 empty] | Cut at C28 | 100 | 20 to 40 | 50 to 200 | t.b.d
| Refractory compounds (organics containing for example S, P, Si) | tbd | PTR-MS [5] *) | 50 | [9 empty] | Breakthrough to be studied for low boiling point S or Si compounds. Cut at C28. | 100 | 20 to 40 | 50 to 200 | t.b.d
|  |  | L-IMS [9] | 800
|  |  | PID [8] | 1000
Reticle Storage (inside stocker, inside pod, inside exposure tool library, inside inspection tool)
| Total Inorganic Acids | <200 | PWPD-IC [2] *) | 10 | IC-Conductivity [13,14] | 10 | 1500-200 | 1-5 | Every 6 months | IC-Conductivity [13,14] | 100 | 7-90 | quarterly
|  |  | SRI-MS [7] *) | 50
|  |  | CRDS [3] *) | 100
| Total Organic Acids [30] | <200 | PWPD-IC [2] *) | 30 | IC-Conductivity [13,14] | 10 | 1500-200 | 1-5 | Every 6 months | IC-Conductivity [13,14] | 100 | 7-90 | quarterly | Use special GC column | 100 | 20 to 40 | 50 to 200 | 2x annual
|  |  | PTR-MS [5] *) | 200
| Total Bases | <200 | PWPD-IC [2] *) | 30 | IC-Conductivity [13,14, 15] | 10 | 1500-200 | 1-5 | Every 6 months | IC-Conductivity [13,14, 15] | 100 | 7-90 | quarterly
|  |  | PAS [6] *) | 20
| Condensable organics (definition to SEMI F21-1102, bp 150 °C) | < 100 | PTR-MS [5] *) | 50 | [9 empty] | Cut at C28. Does not apply to inside POD sampling | 100 | 20 to 40 | 50 to 200 | t.b.d
| Refractory Compounds | tbd | PTR-MS [5] *) | 50 | [9 empty] | Breakthrough to be studied for low boiling point S or Si compounds. Cut at C28. | 100 | 20 to 40 | 50 to 200 | t.b.d
|  |  | L-IMS [9] | 800
Gate/Furnace area wafer environment (FOUP inside)
| Total metals [8] (E+10 atoms/cm2/day) | 0.5 | [16 empty] | Witness wafer + VPD ICP-MS [17] or TXRF [18] (but not for light the elements Li, Na, Al & Mg) or equivalent. SEMI E45-1101. | 1.0E+09 atoms/cm2
| Dopants [4] (E+10 front end of line only) | 0.5 | PTR-MS [5] *) | 50 | [14 empty] | Witness wafer + DSE ICP-MS [19] (TOF-SIMS [20] also or equivalent). SEMI E45-1101. | 1.0E+09 atoms/cm2
| Volatile Organics (w/ GCMS retention times ≥ benzene, calibrated to hexadecane) [31] | 2000 | [11 empty] | Cut at C28. Be aware that the FOUP is not tight and will be replenished with gases from outside the FOUP. Make sure the surrounding environment is clean enough. | 100 | 10 to 20 but always identical volumes, temperature and background | 50 to 200 | t.b.d
| SMC (surface molecular condensable) organics on wafers, ng/cm2/day [12] | tbd | [16 empty] | Witness wafer TD-GC/MS SEMI MF 1982-1103 | 0.1 ng/cm2
Salicidation Wafer Environment (FOUP inside; wafer environment)
| Total Inorganic Acids | 2000 | PWPD-IC [2] *) | 10 | IC/ Conductivity [13, 14] | 100 | 200-500 | 1-5 | quarterly | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
| Total Organic Acids [30] | 5000 | PWPD-IC [2] *) | 30 | IC/ Conductivity [13, 14] | 100 | 200-500 | 1-5 | quarterly | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly | Use special GC column | 100 | 10 to 20 but always identical volumes, temperature and background | 50 to 200 | t.b.d
|  |  | PTR-MS [5] *) | 200
Exposed Copper (POI, Wafer Environment)
| Total Inorganic Acids | 500 | PWPD-IC [2] *) | 10 | IC/ Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
| Total Bases | 2000 | gg | 10 | IC/ Conductivity [13, 14, 15] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14, 15] | 100 | 7-90 | quarterly | For organic amines only (not total bases). Cut at C28. | 100 | 20 to 40 | 50 to 200 | t.b.d
|  |  | PAS [6] *) | 20
|  |  | IMS [1] | 1000
| Total Organic Acids [30] | 500 | PWPD-IC [2] *) | 30 | IC/ Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly | Use special GC column | 100 | 20 to 40 | 50 to 200 | 2x annual
|  |  | PTR-MS [5] *) | 200
| Total other corrosive species [32] | 1000 | IMS [1] | 1000 | IC/ Conductivity [13, 14] | 100/1000 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100/- | 7-90 | quarterly
| H2S | 1000 | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
| Total sulphur compounds | 2500 | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
|  |  | PTR-MS [5] *) | 1000
Exposed Aluminum (POI, Wafer Environment)
| Total Inorganic Acids | 500 | PWPD-IC [2] *) | 10 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
| HCl | 100 | PWPD-IC [2] *) | 10 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
| HF | 200 | PWPD-IC [2] *) | 10 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
| HBr | tbd | PWPD-IC [2] *) | 10 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
| Total Organic Acids [30] | tbd | PWPD-IC [2] *) | 30 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | Every 6 months | IC/ Conductivity [13, 14] | 100 | 7-90 | quarterly | Use special GC column | 100 | 20 to 40 | 50 to 200 | 2x annual
|  |  | PTR-MS [5] *) | 200
| Total other corrosive species [32] | tbd | IMS [1] | 1000 | IC-Conductivity [13, 14] | 100/1000 | 200-500 | 1-5 | Every 6 months | IC-Conductivity [13, 14] | 100/- | 7-90 | quarterly
Exposed Copper Wafer Environment (FOUP inside)
| HCl | 200 | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
| HF | 5000 | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
|  |  | CRDS [3] *) | 100
|  |  | IMS [1] | 1000
| HBr | tbd | IMS [1] | 1000 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
| HNO3 | tbd | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
|  |  | CLD [10] | 200
| Total Organic Acids [30] | 100 | PWPD-IC [2] *) | 400 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed | Use special GC column | 100 | 10 to 20 but always identical volumes, temperature and background | 50 to 200 | as needed
| Total bases | tbd | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14, 15] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14, 15] | 100 | 7-90 | as needed | Cut at C28. Be aware that the FOUP is not tight and will be replenished with gases from outside the FOUP. Make sure the surrounding environment is clean enough. | 100 | 10 to 20 but always identical volumes, temperature and background | 50 to 200 | t.b.d
|  |  | PAS [6] *) | 20
|  |  | IMS [1] | 1000
|  |  | CRDS [3] *) | 4000
| Total other corrosive species [32] | tbd | IMS [1] | 1000 | IC-Conductivity [13, 14] | 100/1000 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
| H2S | tbd | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
| Total sulphur compounds | 1000 | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
|  |  | PTR-MS [5] *) | 1000
Exposed Aluminum Wafer Environment (FOUP inside)
| Total Inorganic Acids | tbd | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13,14] | 100 | 7-90 | as needed
|  |  | CRDS [3] *) | 100
|  |  | UV-F [4] *) | 500
|  |  | IMS [1] | 1000
| HCl | 100 | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
| HF | 200 | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
|  |  | CRDS [3] *) | 100
| HBr | tbd | IMS [1] | 1000 | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
| HNO3 | tbd | PWPD-IC [2] *) | 200 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
|  |  | CLD [10] | 200
| Total Organic Acids [30] | tbd | PWPD-IC [2] *) | 400 **) | IC-Conductivity [13, 14] | 100 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed | Use special GC column | 100 | 10 to 20 but always identical volumes, temperature and background | 50 to 200 | as needed
|  |  | PTR-MS [5] *) | 200
| Total other corrosive species [32] | tbd | IMS [1] | 1000 | IC-Conductivity [13, 14] | 100/1000 | 200-500 | 1-5 | as needed | IC-Conductivity [13, 14] | 100 | 7-90 | as needed
Methods and instruments are known and are being optimized
Methods and instruments are known and in preparation for production
|  | Methods and instruments are in beta tool status under qualification | [5 empty] | ¿
Footnotes from first row : Please refer to YE3 table footnotes.
[23] note amendment: entry to scanner has two applications: supply gas with 10 ppt LDL and filtered ambient air with 100 ppt LDL.
Footnotes
[1] IMS Ion Mobility Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[2] PWPD-IC Parallel Wet Plate Denuder Ion Chromatography; for method key features and capabilities please refer to the methods overview table (YE 4a)
[3] CRDS Cavity Ring Down Spectroscopy; for method key features and capabilities please refer to the methods overview table (YE 4a)
[4] UV-F UV Fluorescence; for method key features and capabilities please refer to the methods overview table (YE 4a
[5] PTR-MS Proton Transfer Mass Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[6] PAS Photoaccoustic Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[7] SRI-MS Switchable Reactive Ion MS (matrix nitrogen or inerts); for method key features and capabilities please refer to the methods overview table (YE 4a)
[8] PID Photoionization Detektor; for method key features and capabilities please refer to the methods overview table (YE 4a)
[9] L-IMS Laser Induced Ion Mobility Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[10] CLD Chemiluminescence Detection of NO y compounds, for method key features and capabilities please refer to the methods overview table (YE 4a)
*) Method is capable to detect and quantity individual compounds. Sum values are therefore the results of summing up individual concentrations
**) LDL calculated from LDL of unrestricted volume & sampling time (clean room) to restricted volume & sampling time (FOUP)
[11] Impinger sampling for FOUP and other micro-environments is only applicable for continuous purge outgas testing, not for static volume testing, due to the limited volume of the device.
[12] Due to its long sampling time, passive sampling (formerly beaker sampling) is recommended only for long-term averaging of AMC concentrations.
[13] Ion Chromatography Mass Spectrometry (IC/MS) may be used to identify unknown species
[14] Capillary electrophoresis mass spectrometry (CE/MS) may be used with probable higher detection limits
[15] Liquid chromatography mass spectrometry (LC/MS) may be used for organic bases
[16] Open beakers are not suitable for compressed gas analysis. Dry passive samplers would work.
[17] VPD ICP-MS: Vapor phase deposition inductively coupled plasma mass spectrometry
[18] TXRF: Total reflectance X ray spectroscopy
[19] DSE ICP-MS: Drop scan extraction inductively coupled plasma mass spectrometry
[20] TOF-SIMS: Time of flight secondary ion mass spectrometry

## Notes for Table YE4
INDEX | Notes for Table YE4 AMC Monitoring
(Footnotes from first row : Please refer to YE3 table footnotes.)
[23] note amendment: entry to scanner has two applications: supply gas with 10 ppt LDL and filtered ambient air with 100 ppt LDL.
Footnotes
[1] IMS Ion Mobility Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[2] PWPD-IC Parallel Wet Plate Denuder Ion Chromatography; for method key features and capabilities please refer to the methods overview table (YE 4a)
[3] CRDS Cavity Ring Down Spectroscopy; for method key features and capabilities please refer to the methods overview table (YE 4a)
[4] UV-F UV Fluorescence; for method key features and capabilities please refer to the methods overview table (YE 4a)
[5] PTR-MS Proton Transfer Mass Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[6] PAS Photoaccoustic Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[7] SRI-MS Switchable Reactive Ion MS (matrix nitrogen or inerts); for method key features and capabilities please refer to the methods overview table (YE 4a)
[8] PID Photoionization Detektor; for method key features and capabilities please refer to the methods overview table (YE 4a)
[9] L-IMS Laser Induced Ion Mobility Spectrometry; for method key features and capabilities please refer to the methods overview table (YE 4a)
[10] CLD Chemiluminescence Detection of NO y compounds, for method key features and capabilities please refer to the methods overview table (YE 4a)
*) Method is capable to detect and quantity individual compounds. Sum values are therefore the results of summing up individual concentrations
**) LDL calculated from LDL of unrestricted volume & sampling time (clean room) to restricted volume & sampling time (FOUP)
[11] Impinger sampling for FOUP and other micro-environments is only applicable for continuous purge outgas testing, not for static volume testing, due to the limited volume of the device.
[12] Due to its long sampling time, passive sampling (formerly beaker sampling) is recommended only for long-term averaging of AMC concentrations.
[13] Ion Chromatography Mass Spectrometry (IC/MS) may be used to identify unknown species
[14] Capillary electrophoresis mass spectrometry (CE/MS) may be used with probable higher detection limits
[15] Liquid chromatography mass spectrometry (LC/MS) may be used for organic bases
[16] Open beakers are not suitable for compressed gas analysis. Dry passive samplers would work.
[17] VPD ICP-MS: Vapor phase deposition inductively coupled plasma mass spectrometry
[18] TXRF: Total reflectance X ray spectroscopy
[19] DSE ICP-MS: Drop scan extraction inductively coupled plasma mass spectrometry
[20] TOF-SIMS: Time of flight secondary ion mass spectrometry

## Table YE4a
INDEX
Table YE 4a Supporting Table for On-line Methods
| Class and type of Molecular Contamination |  | Limit of quantification classified | Measurement Range | Matrix | Data interval | Sample flow** | response time t95 | recovery time t05 | zero point*** | span drift**** | interference | service interval | multipoint | measurement | literature reference
| (to SEMI F 21-1102, ITRS Definitions of AMC and CAS-N°) | CAS-N° | Class 1 (10 pptv) to Class 5 (100'000 pptv) | (pptv)* | (air, others) | (sec) | (ml/min) | rise (sec) | decline (sec) | drift (%/month) | (%/month) | to (% of signal) | (months) | sampling capability***** | principle | ( publication with IDM/academia)
MA (Molecular Contamination of Acid or acid precursor type)
| Total Inorganic Acids as sum value****** |  | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% | 5% of NOx | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
| HF (Hydrogen Fluoride) as individual compound | 7664-39-3 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  | 7664-39-3 | Class 2 (30 pptv) | 0 - 10E+6 pptv | air | 5 | < 600 | < 600 | < 600 | 12 months | 0.5 ppb/month |  | 12 months | up to 16 ports | Cavity Ring-Down Spectroscopy (CRDS) | John C. Bosch Jr, dec 2010 - US EPA
|  | 7664-39-3 | Class 2 (100 pptv) | 0-10E+8 pptv | air & inerts | 1 | 1000 |  |  | 0 |  | 5% of HCl | no service required | yes | Continuous Wave - CRDS
|  | 7664-39-3 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% | 5% of NOx | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 7664-39-3 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
| HCl (Hydrogen Chloride) as individual compound | 7647-01-0 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  | 7647-01-0 | Class 3 (500pptv) | 0-2.5E+9 pptv | air & inerts | 1 | 1000 |  |  | 0 |  |  | no service required | Yes | Continuous Wave - CRDS
|  | 7647-01-0 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% | 5% of NOx | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 7647-01-0 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
| HBr (Hydrogen Bromide) as individual compound | 10035-10-6 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% | 5% of NOx | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 10035-10-6 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
| HNO2 (nitrous acid) as individual compound | 7782-77-6 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
| HNO3 (nitric acid) as individual compound | 7697-37-2 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
| NO y (representing HNO3 + HNO2) |  | Class 2 (100 pptv) | 0-1E+6 pptv | air | 10 | 300 | 3 | 10 | <1% | <1% | CO2 | 12 months | yes | Chemiluminescence CLD
| NOx (representing NO+NO2) |  | Class 3 (1000 pptv) | 0-4E+5 pptv | air | 30 | 250 | < 1 | < 20 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
| SO2 (Sulfur Dioxide) as individual compound | 7446-09-5 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  | 7446-09-5 | Class 2 (20 pptv) |  | nitrogen | var. | 50-1000 | 1 | 0.2 | none |  |  | 6 months | yes | Switchable Reactive Ion MS (SRI-MS)
|  | 7446-09-5 | Class 2 (100 pptv) | 0-1E+5 pptv | Air | 30 | 250 | < 1 | < 20 | 0% (autozero) | 5% | 5% of NOx | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 7446-09-5 | Class 3 (500 pptv) | 0-5E+5 pptv | air | 5 | 700 | 120 | 80 | <0,1% | <0,1% | 2% of NOx | 3 months |  | UV Fluorescence (UVF)
|  | 7446-09-5 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
Others (as individual compounds)
Total Organic Acids as sum value******
| HFc (Formic Acid) as individual compound | 64-18-6 | Class 2 (30 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  | 64-18-6 | Class 3 (500 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 64-18-6 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
| HAc (Acetic Acid) as individual compound | 64-19-7 | Class 2 (30 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  | 64-19-7 | Class 3 (500 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 64-19-7 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
| HPr (Propionic Acid) as individual compound | 79-09-4 | Class 6 (200 000 pptv) | 0-5E+7 pptv | air | [9 empty] | Fourier-Transform IR (FT-IR)
|  | 79-09-4 | Class 3 (500 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| HLac (Lactic Acid) as individual compound | 50-21-5 | Class 3 (500 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| Hox (Oxalic Acid) as individual compound | 144-62-7 | Class 3 (500 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
Others (as individual compounds)
MB (Molecular Contamination of alkaline nature, bases)
| Total Base as sum value****** |  | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
| NH3 (Ammonia) as individual compound | 7664-41-7 | Class 2 (20 pptv) | 0-1E+6 pptv | air | 15 | 1000 | 60 | 60 | < 2% | < 2% | water vapour | 12 months | yes | Photoaccoustic Spectrometry (PAS) | Besson et al. (2006)
|  | 7664-41-7 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  | 7664-41-7 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 7664-41-7 | Class 4 (4000 pptv) | 0-400E+6 pptv | air & inerts | 1 | 1000 | [5 empty] | no service | yes | Continuous Wave - CRDS
| TMAH (Trimethylammoniumhydroxide) as individual compound | 75-59-2 | Class 2 (100 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 75-59-2 | Class 3(700 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| Me3N (Trimethylamine) as individual compound | 75-50-3
| Other amines as individual compounds, MEA, N-MEA |  | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et. Al. (2011)
|  |  | Class 3 (1000 pptv) | 0-1E+6 pptv | air | 5 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  |  | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
| NMP (N-Methylpyrolidone) as individual compound | 872-50-4 | Class 1 (10 pptv) | 0-1E+6 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 872-50-4 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 872-50-4 | Class 3 (1000 pptv) | 0-1E+6 pptv | air | 5 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| Total reduced Nitrogen, total amine****** |  | Class 3 (500 pptv) | 0-1E+6 pptv | air | 30 | 300 | 60 | 60 | <2% | <2% | CO2 | 12 months | yes | Chemiluminescence CLD
Other airborne compounds or alkaline nature
Molecular Contamination with Corrosive effects
Total Corrosive Species as sum value******
| F2 (Fluorine) as individual value | 7782-41-4
| Cl2 (Chlorine) as individual value | 7782-50-5 | Class 3 (1000 pptv) | 0-4E+5 pptv | air | 30 | 250 | < 1 | < 20 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
| I2 (Iodine) as individual value | 7553-56-2
| Ozone |  | Class 3 (500 pptv) | 0-1E+6 pptv | air | 5 | 700 | 75 | 50 | <0,1% | <0,1% |  | 3 months | yes | UV Absorption (UV-A)
other corrosive species as individual value
MC (Molecular Contamination of Condensable Nature, Organics)
Condensable Organics to SEMI F-21-95 (bp > 150 °C), sum value******
| Volatile Organic Compounds (VOC) retention time > benzene |  | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 120 | 250 | < 10 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Photoionizsation Detector (PID)
| PGMEA (Methoxy-propyl-acetate) as individual compound | 108-65-6 | Class 2 (100 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 108-65-6 | Class 3(700 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 108-65-6 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Photoionizsation Detector (PID)
| EL (Ethyl lactate) as individual compound | 97-64-3 | Class 2 (100 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 97-64-3 | Class 3(600 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 97-64-3 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Photoionizsation Detector (PID)
Other organic compounds (as individual compounds)
Molecular Contamination of refractive type
Total Refractory compounds (containing Si, P, S) as sum value******
| HMDS (Hexamethyl-di-silazane) as individual compound | 999-97-4 | Class 2(50 pptv) | 0-10E+8 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 999-97-4 | Class 3(800 pptv) | 0-1E+5 pptv | air | 30 | 200 | 5 | 10 |  |  |  | 3 months | yes | Laser Induced IMS (L-IMS)
|  | 999-97-3 | Class 3(1000 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 999-97-4 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 | < 180 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Photoionizsation Detector (PID)
| TMS (Trimethyl-silanol) as individual compound | 1066-40-6 | Class 2(50 pptv) | 0-10E+8 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 1066-40-6 | Class 3(1000 pptv) | 0-10E+7 pptv | air | 5 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| HMDSO (Hexa-methyl-di-siloxane) as individual compound | 107-46-0 | Class 2(50 pptv) | 0-10E+8 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 107-46-0 | Class 3(1000 pptv) | 0-10E+7 pptv | air | 5 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  | 107-46-0 | Class 3 (1100 pptv) | 0-10E+5 pptv | air | 30 | 200 | 5 | 10 |  |  |  | 3 months | yes | Laser Induced IMS (L-IMS)
| other linear siloxane compounds (as individual compounds) |  | Class 2(50 pptv) | 0-10E+8 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  |  | Class 3(1000 pptv) | 0-10E+7 pptv | air | 2 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| other cyclic siloxane compounds (as individual compounds) |  | Class 2(50 pptv) | 0-10E+8 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  |  | Class 3(1000 pptv) | 0-10E+7 pptv | air | 2 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| organo-phosphorous compounds (as individual compounds) |  | Class 2(50 pptv) | 0-10E+8 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
|  |  | Class 3(1000 pptv) | 0-10E+7 pptv | air | 2 | 50 | 5 | 10 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
organo-sulfur compounds (as individual compounds)
MD (Molecular Contamination of Dopant type)
Total Dopant compounds (n-dopants or p-dopants)******
Boron compounds (airborne, of inorganic nature)
Boron compounds (airborne, of organic nature)
| Phosphorous compounds (airborne, of inorganic nature), PH3 | 7802-51-2 | Class 5 (100'000 pptv) | 0-5E+6 pptv | air | 30 | 250 | < 1 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Photoionizsation Detector (PID)
Phosphorous compounds (airborne, of organic nature)
| Arsenic compounds (airborne, or inorganic nature), AsH3 | 7784-42-1 | Class 5 (100'000 pptv) | 0-5E+6 pptv | air | 30 | 250 | < 1 | < 60 | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Photoionizsation Detector (PID)
Arsenic compounds (airborne, or organic nature)
Other dopant compounds (as individual compound)
Molecular Contamination of Sulfurous Nature
Total Sulfurous Compounds as sum value******
| H2S (hydrogen sulfide) as individual compound | 7783-06-4 | Class 3 (500 pptv) | 0-5E+5 pptv | air | 20 | 700 | 120 | 80 sec | <0,1%/month | <0,1%/month |  | 3 months | yes | Converter/UV Fluorescence (UVF)
|  | 7783-06-4 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 |  | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 7783-06-4 | Class 4 (10 000 pptv) | 0-5E+5 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Selective Reactive Ion MS (SRI-MS) | A. Feilberg et.al. (2010)
|  | 7783-06-4 | Class 5 (100 000 pptv) | 0-5E+8 pptv | air & inerts | 1 | 1000 |  |  | 0 |  |  | no service required | Yes | Continuous Wave - CRDS
| SO2 (Sulfur Dioxide) as individual compound | 7446-09-5 | Class 1 (10 pptv) |  | air & inerts | 900 | 5000 | 900 | 900 |  |  |  | 1 month | yes | Parallel-Wet-Plate Denuder/IC (PWPD-IC) | M. Takeuchi et.al. (2011)
|  | 7446-09-5 | Class 2 (100 pptv) | 0-1E+5 pptv | air | 30 | 250 | < 1 | < 20 | 0% (autozero) | 5% | 5% of NOx | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
|  | 7446-09-5 | Class 2 (20 pptv) |  | nitrogen | 1 | 50-1000 | 0.5 | 1 | none |  |  | 3 months | yes | Switchable Reactive Ion MS (SRI-MS)
|  | 7446-09-5 | Class 3 (500 pptv) | 0-5E+5 pptv | air | 5 | 700 | 120 | 80 | <0,1%/month | <0,1%/month | 2% of NOx | 3 months | yes | UV Fluorescence (UVF)
| Mercaptanes as individual compounds |  | Class 3(1000 pptv) | 0-10E+7 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS) | D. Jaksch et.al. (2004)
| DMS (Dimethyl-sulfide) as individual compound | 75-18-3 | Class 3(1000 pptv) | 0-10E+7 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS) | Warnecke et.al. (2003)
|  | 75-18-3 | Class 3 (1000 pptv) | 0-5E+5 pptv | air | 30 | 250 | < 1 |  | 0% (autozero) | 5% |  | 3-6 months | up to 64 ports | Ion Mobility Spectrometry (IMS)
| DMSO (Dimethly-sulfoxide) as individual compound | 67-68-5 | Class 3(1000 pptv) | 0-10E+7 pptv | air | 2 | 50 | 1 | 2 |  |  |  | 3 months | yes | Proton Transfer MS (PTR-MS)
| * Range that can be calibrated and renders reliable results | [13 empty] | Methods and instruments are known and are being optimized
| ** Indicate also typical sampled volume to render a datapoint/reading | [13 empty] | Methods and instruments are known and in preparation for production
| *** Indicate typical time period for drift consideration and zero point re-calibration | [13 empty] | Methods and instruments are in beta tool status under qualification
**** Indicate typical time period for span drift consideration and re-calibration
***** Monitor can be used in an automated multi-sampling point environment; how many ports/time resolution?
****** "Total" analyzers may have different response factors to individual contaminants, and therefore may over or under report actual gas-phase concentrations.
Besson et. al. (2006), Ammonia trace measurments at ppbv level based on near-IR photoaccoustic spectroscopy, Appl. Phys. B (2006)
Feilberg et. Al (2010), Odorant Emissions from intensive pig production measured by on-line PTR-Mass Spectrometry, Environ. Sci. Technol. (2010), 44, 5894-5900
Jaksch et. Al. (2004), The effect of ozone treatment [ ] using PTR-MS, Int. Journal of Mass Spectrometry 239 (2004) 209-214
Takeuchi et.Al. (2011), Parallel-Plate Wet Denuder Coupled Ion Chromatograph for Near-Real-Time Detection of Trace Acid Gases in Clean Room Air, Proceedings on ISSM (2010), IUC-P-058I
Warnecke et. Al (2003), Validation of atmospheric VOC measurments by PTR-MS using a ga-chromatographic pre-separation method, Environ. Sci. Technol. (2003), 37, 2494-2501

## Notes for Table YE4a
INDEX
Notes for Table YE4a
* Range that can be calibrated and renders reliable results
** Indicate also typical sampled volume to render a datapoint/reading
*** Indicate typical time period for drift consideration and zero point re-calibration
**** Indicate typical time period for span drift consideration and re-calibration
***** Monitor can be used in an automated multi-sampling point environment; how many ports/time resolution?
****** "Total" analyzers may have different response factors to individual contaminants, and therefore may over or under report actual gas-phase concentrations.
Besson et. al. (2006), Ammonia trace measurments at ppbv level based on near-IR photoaccoustic spectroscopy, Appl. Phys. B (2006)
Feilberg et. Al (2010), Odorant Emissions from intensive pig production measured by on-line PTR-Mass Spectrometry, Environ. Sci. Technol. (2010), 44, 5894-5900
Jaksch et. Al. (2004), The effect of ozone treatment [ ] using PTR-MS, Int. Journal of Mass Spectrometry 239 (2004) 209-214
Takeuchi et.Al. (2011), Parallel-Plate Wet Denuder Coupled Ion Chromatograph for Near-Real-Time Detection of Trace Acid Gases in Clean Room Air, Proceedings on ISSM (2010), IUC-P-058I
Warnecke et. Al (2003), Validation of atmospheric VOC measurments by PTR-MS using a ga-chromatographic pre-separation method, Environ. Sci. Technol. (2003), 37, 2494-2501

## Table YE5
INDEX | NEW: Table YE5 Yield requirements for MEMS production
| Year of Production | 2017 | Mid term
Infrared inspection with automated and fullwafer scan
| Defect types for surface MEMS technology(A) | particles, foreign materials, cracks
| Defect types for bulk technology(A) | wafer bond quality, buried parictles, cracks
| Wafer thickness for r surface MEMS technology(A) | 250 µm–1,5 mm
| Wafer thickness for for bulk technology(A) | 300 µm–4 mm
| Thruput at Resolution 4 µm (200 mm wa/hrs | 25
| Thruput as Resolution 1 µm (200 mm wa/hrs) | 5
| Illumination mode | reflexion and transmission
| Wafer size (mm) | 200 | 300
| Wafer bow plus/minus for handling and automatic focus adjustment (µm) | 400 | 600
Review tool for MEMS wafer with backside structure
| Wafer thickness for surface MEMS technology(A) | 250 µm–1,5 mm
| Wafer thickness for bulk technology(A) | 300 µm–4 mm
| Wafer backside handling with wafer edge grip (mm) | 5
| Wafer bow plus/minus for handling and automatic focus adjustment (µm) | 400 | 600
| Thruput (at 50 defects per wafer) | 100 wa/hrs
| Illumination mode | brightfield, darkfield
| Additional illumination mode for bulk technology | infrared
CD Sem under angle of 45
| Tilt | 0–60
| Min, Structure | 1 µm | 300 nm
| Max Structure, larger structures can be measured with optical methods | 10 µm
| Aspect ratio | 10–30 µm depth
| Resolution | tbd
| Thruput | tbd
| Precision | tbd
| Repeatability | tbd

## Notes for Table YE5
INDEX
Notes for Tables YE5:
[A] definition of technology: till 30 µm depth = surface technology, stacked technologies and those with depth more than 30 µm = bulk technology

## Figure YE5
INDEX | Wafer Environment and Reticle Environment Contamination Measurement and Control Potential Solutions
|  | SUBSTRATE |  | SUBSTRATE LOCATION
| First Year of IC Production DRAM 1/2 Pitch | WAFER | RETICLE | STORAGE | PROCESS | TRANSPORT | 2017 20nm | 2018 18nm | 2019 16nm | 2020 14nm | 2021 | 2022
New FOUP Outgassing / Contamination
| Off line AMC measurement in laboratory | X | 0 | X | 0 | 0
| Inline AMC measurement inside FOUP in fab | X | 0 | X | X | X
| Inline AMC measurement inside 450mm FOUP and MAC in fab | X | 0 | X | X | X
Used FOUP Outgassing / Contamination (no wafer inside)
| Off line AMC measurement in laboratory | X | 0 | X | 0 | 0
| Inline AMC measurement inside FOUP in fab | X | 0 | X | X | X
| Inline AMC measurement inside 450mm FOUP and MAC in fab | X | 0 | X | X | X
| Vacuum purge and heating | X | 0 | X | 0 | 0
Wafer Cross Contamination Inside FOUP (e.g acids after dry etch, oxygen or humidity before EPI clean)
| Inline AMC measurement inside FOUP in fab | X | 0 | X | X | X
| Integrated AMC measurement inside load port | X | 0 | 0 | X | 0
| Vacuum purge : outgass the FOUP with wafers under vacuum and fill it with N2 | X | 0 | X | 0 | X
| N2 purge station : injection of N2 inside FOUP with wafer | X | 0 | X | 0 | X
| N2 purge station integrated in stocker : injection of N2 inside FOUP with wafers | X | 0 | X | 0 | 0
| FOUP change : wafer transfer in a clean FOUP during q-time | X | 0 | X | 0 | X
| Purgeable load port : injection of N2 when the FOUP is connected to the EFEM | X | 0 | 0 | X | 0
| Outgassing chamber integrated in process equipment | X | 0 | 0 | X | 0
| Wafer transfer under vacuum | X | 0 | 0 | X | X
Wafer environment (outside of pod)
| Off line AMC measurement in laboratory | X | 0 | X | 0 | X
| Inline AMC measurement of WE | X | 0 | X | X | X
This legend indicates the time during which research, development, and qualification/pre-production should be taking place for the solution.
Research Required
Development Underway
Qualification / Pre-Production
Continuous Improvement
Work-in-Progress--Do Not Publish

## Figure YE6
INDEX | Substrate Environment Contamination Control Potential Solutions- Critical Components
| First Year of IC Production | 2017 | 2018 | 2019 | 2020 | 2021
| Logic 1/2 pitch, nm (contacted) | 18 | 18 | 14 | 14 | 12
GENERAL
Fluid purity impact on device yield / performance
Contaminant based process control
Critical Components
Particle quality analysis of the critical components, 10 nm [A]
Drivers:
[A] killer particle control
[B] critical particle effect
[C] critical particles and organics
[D] critical organics
This legend indicates the time during which research, development, and qualification/pre-production should be taking place for the solution.
Research Required
Development Underway
Qualification/Pre-Production
Continuous Improvement

## Figure YE7
INDEX | Substrate Environment Contamination Control Potential Solutions - UPW
| First Year of IC Production | 2017 | 2018 | 2019 | 2020 | 2021
| Logic 1/2 pitch, nm (contacted) | 18 | 18 | 14 | 14 | 12
GENERAL
Fluid purity impact on device yield / performance
Contaminant based process control
Ultra Pure Water
Particle online metrology 10 nm [A]
Particle online metrology 5 nm [A]
Effective Filtration at 10nm particle size [A]
Effective Filtration at 5nm particle size [A]
Improved components to reduce particle generation (IX resin, coated pumps, etc.) [C]
Off-line test for particle filtration performance validation [A]
Mitigation measures to control EAP (electrically active particles) [A]
Mitigation measures to control High Molecular Weight Polymers [C]
On-line detection of the critical organics (non-volatile) needs development [D]
Use on-line TOC as way to control total organics (assuming that all is critical) [D]
Need to define effect of UPW borne H2O2 to manufacturing process
Need to better define effect of Metal in UPW to image sensors
Drivers:
[A] killer particle control
[B] critical particle effect
[C] critical particles and organics
[D] critical organics
This legend indicates the time during which research, development, and qualification/pre-production should be taking place for the solution.
Research Required
Development Underway
Qualification/Pre-Production
Continuous Improvement

## Figure YE8
INDEX | Substrate Environment Contamination Control Potential Solutions - Liquid Chemicals
| First Year of IC Production | 2017 | 2018 | 2019 | 2020 | 2021
| Logic 1/2 pitch, nm (contacted) | 18 | 18 | 14 | 14 | 12
GENERAL
Fluid purity impact on device yield / performance. Requirements are specific to the chemistry and the Process step that it is used for.
Contaminant based process control ???
CHEMICALS - Metrology
Development of 40 nm particle counter [A]
Development of 20 nm particle counter [A]
Development of 10 nm particle counter [A]
Particle characterization to identify source of contamination [A], (for components and systems)
Micelle particle size and concentration characterization [B] for surfactant containing chemistries.
Bubble discrimination of reactive chemistries [A]
Improved metrology for concentration measurements [B]
Anion & cation measurement in process chemicals, e.g. cleaning chemicals [A]
Organic measurement (TOC and speciation) in process chemicals, e.g. H2O2, IPA [A]
Contaminant characterization in CMP slurries, e.g. zeta potential, large particle size [A]*
Contaminant characterization in CMP slurries. (Agglomertaes, Foreign Material) [B]
Characterization of CMP slurries, e.g. Particle size distribution [B]
On-line contaminant and constituent measurement in plating chemicals, e.g. copper sulfate, organic additives [A],[B]
Improved blend accuracy taking into account incoming chemical concentration changes (closed loop control capability) [B]
Closed loop point of use blending improvements [B]
CHEMICALS - Components and Systems
Filtration 0.02/0.04 µm with high flux (1 gpm/0.5 psi/10"/1cP) [A]
Filtration sub 0.02 µm with high flux (1 gpm/0.5 psi/10"/1cP) [A]
Filtration sub 0.01 µm with high flux (1 gpm/0.5 psi/10"/1cP) [A]
Ionic POU purifiers (<1 ppt metals) [A]
Inline End point detection for Ionic POU purifiers [A]
Improved components to reduce particle generation, e.g. tubing, valves, pumps, fittings, etc. [A]
Improved system startup & flushing methodology to more quickly reach particle target and reduce chemical usage [A]
Improved flow capacity and increased pressure for pumps [C]
Improved flow capacity increased pressure and temperature for piping components [C]
Reduced pressure fluctuations by improved BCD system performance and by better pressure and flow control [A], [B]
Improved connection technology (reduce leakages and increase reliabilitiy) [A], [D]
Decrease permeation of piping components (eg HCl permeation) [A], [D]
Higher purity resins and components (metals, organics, anions, cations, surface smoothness, permeability) [A]
Elimination of particle sources in chemical distribution system including POU (particle prevention in chemical conditions) [A], (for EUV mask cleaning)
*Slurries are also looked at in terms of changes in active components (e.g., H2O2)
Driver:
[A] Purity
[A] Purity - components and systems
[B] Process control
[C] Capacity
[D] Maintenance
This legend indicates the time during which research, development, and qualification/pre-production should be taking place for the solution.
Research Required
Development Underway
Qualification/Pre-Production
Continuous Improvement